| THE MINISTRY OF HEALTH OF VIETNAM ------- | THE SOCIALIST REPUBLIC OF VIET NAM Independence-Freedom-Happiness ----------------- |
| --- | --- |
| No. 04/VBHN-BYT | Hanoi, June 03, 2024 |

DECREE

PRESCRIBING MEDICAL DEVICE MANAGEMENT [1]

The Government’s Decree No. 98/2021/ND-CP dated November 08, 2021 prescribing medical device management[2], coming into force from January 01, 2022, is amended by:

The Government’s Decree No. 07/2023/ND-CP dated March 03, 2023 providing amendments to the Government’s Decree No. 98/2021/ND-CP dated November 08, 2021 prescribing medical device management[3], coming into force from March 03, 2023.

The Government’s Decree No. 96/2023/ND-CP dated December 30, 2023 on elaboration of the Law on Medical Examination and Treatment, coming into force from January 01, 2024.

Pursuant to the Law on Government Organization dated June 19, 2015; the Law on Amendments to the Law on Government Organization and the Law on Local Government Organization dated November 22, 2019;

Pursuant to the Law on Investment dated June 17, 2020;

At the request of the Minister of Health of Vietnam;

The Government promulgates a Decree prescribing medical device management. [4]

Chapter I

GENERAL PROVISIONS

Article 1. Scope

1. This Decree deals with the management of medical devices [5], including: classification of medical devices[6]; manufacture, placement on the market, trading, import, export and provision of services related to medical devices[7]; medical device information and advertising[8]; management of prices of medical devices, and management[9] and use of medical devices[10] in health facilities.

2. This Decree does not apply to:

a) Ingredients and semi-finished products used for manufacture of medical devices[11], except raw materials containing narcotic substances and precursors;

b) Raw materials for manufacture of medical devices [12], including samples of blood, serum, plasma, urine, fecal, body fluids or other human materials that are subject to biosafety requirements upon import/export as prescribed by law;

c) Medical gas;

d) Medical device accessories[13];

dd) Research Use Only (RUO) and Laboratory Use Only (LUO) products.

Article 2. Definitions

1. “medical device” [14]  means any instrument, implant, apparatus, material, in-vitro reagent or calibrator, or software that meets all of the following requirements:

a) It is intended, by the product owner [15], to be used, whether alone or in combination, for human beings for the purpose of one or more of the following:

- diagnosis, prevention, monitoring, treatment or alleviation of disease, or compensation for an injury or trauma;

- investigation, replacement, modification or support of the anatomy or of a physiological process;

- supporting or sustaining life;

- control of conception,

- disinfection of medical devices [16];

- providing information serving diagnosis, monitoring or treatment through examination of specimens derived from the human body.

b) The device does not achieve its principal intended action in or on the human body by pharmacological, immunological or metabolic means, but that may be assisted in its function by such means to serve the purposes mentioned in Point a of this Clause.

2. “in vitro diagnostic (IVD) medical device” [17] means a reagent, calibrator, control material, instrument, apparatus, equipment or system or other product, whether used alone or in combination, intended by the product owner, to be used in vitro for the examination of specimens derived from the human body.

3. “personalized medical device” [18] means a medical device[19] that is specifically made in accordance with a duly qualified medical practitioner’s written prescription, which gives, under his responsibility, specific design characteristics and intended for the sole use of a particular individual.

4. “accessory” means an article that is intended specifically by its product owner[20] to be used together with a particular medical device[21] to enable or assist that device to be used in accordance with its intended purpose.

5. “product owner” [22] means any organization or person that:

a) supplies the medical device[23] under its/his own name, or under any trademark, design, trade name or other name or mark owned or controlled by it/him; and

b) is responsible for designing, manufacturing, assembling, processing, labeling, packaging, refurbishing or modifying the medical device[24], or for assigning to it[25] a purpose.

Article 3. Principles of management of medical devices [26]

1. The quality, safety and efficacy of the medical device[27] must be verified.

2. Adequate, accurate and timely information about specifications and intended purpose of the medical device [28] and potential risks to the user must be provided.

3. The traceability of medical devices[29] must be ensured.

4. Medical devices[30] shall be managed in accordance with risk classification rules and relevant national technical regulations and/or national standards issued or recognized by competent regulatory authorities or announced standards applied by organizations/individuals in accordance with regulations of law.

5. Medical devices[31] that are measuring devices or radiation equipment must be managed in accordance with regulations of the law on measurement, the law on atomic energy and regulations herein.

6. Chemicals and preparations intended solely for disinfection of medical devices[32] shall be managed in accordance with regulations herein. Chemicals and preparations intended for other purposes, in addition to the disinfection of medical devices[33], shall be managed in accordance with regulations of law on insecticidal and germicidal chemicals and preparations for medical and household use.

7. Medical devices[34], raw materials or ingredients used for manufacturing of medical devices[35], and substances for external quality assessment that contain narcotic substances and precursors must be managed in accordance with regulations of law on drug control, if they are imported or exported, and regulations herein.

8. Regulations on classification, issuance of registration number, and eligibility requirements for trading of medical devices laid down herein shall not apply to:

a) Software used for medical devices[36];

b) Medical devices[37] that are traded as normal goods and imported as gifts or presents given to individuals or organizations other than health facilities.

Chapter II

CLASSIFICATION OF MEDICAL DEVICES [38]

Article 4. Classes of medical devices [39]

Medical devices[40] shall be classified into the following 4 classes according to their levels of potential risks related to their designs and manufacture[41]:

1. Class-A medical devices[42]: medical devices[43] with low risk.

2. Class-B medical devices[44]: medical devices[45] with low to medium risk.

3. Class-C medical devices[46]: medical devices[47] with medium to high risk.

4. Class-D[48]: medical devices[49] with the highest possible risk.

Article 5. Classification rules for medical devices [50]

1. The classification of medical devices[51] must be carried out in accordance with risk classification rules.

2. In the event that a medical device[52] only has one intended purpose which is assigned into two or more risk levels, that medical device shall be assigned into the class representing the highest risk level.

3. In the event that a medical device[53] has multiple intended purposes and each of which represents various risk levels, that medical device shall be assigned into the class representing the highest risk level.

4. In the event that a medical device[54] is designed to be used in combination with another medical device[55], each of the medical devices[56] shall be classified separately.

IVD medical devices[57] that are equipment or systems used in testing process and reagents, titrants, calibrators or control materials may be assigned into various risk levels but assigned into the class representing the highest level of risks posed by the final purpose of that medical device[58] as a whole. IVD medical devices[59] that are other products used in or supporting testing process may be separately assigned into risk levels.

5. The Minister of Health of Vietnam shall provide detailed guidelines on classification of medical devices[60] in accordance with ASEAN’s treaties on classification of medical devices[61] to which Vietnam is a signatory.

6. The classification of medical devices[62] must be carried out by the classification body whose name is specified in the declaration of applied standards or certificate of registration of medical device.

Article 6. Cancellation of classification results of medical devices [63]

1. Classification results of medical devices[64] shall be cancelled in the following cases:

a) The incorrect classification result reduces the level of risk of the medical device[65];

b) The record of classification results is found fraudulent.

2. Procedures for cancellation:

a) Within 01 working day from the day on which a conclusion mentioned in Clause 1 of this Article is given, the Ministry of Health shall issue a decision to cancel classification result which request the classification body to implement remedial measures (if any) and remove the cancelled medical device [66] classification result from the Portal on management of medical devices[67].

The decision to cancel classification result shall be sent to the medical device[68] classification body, Provincial Departments of Health, General Department of Customs and border checkpoint customs authorities, and published on the Portal on management of medical devices[69].

b) After receiving the decision to cancel classification result, the classification body shall cancel all classification results specified in the decision and implement remedial measures against its violations (if any).

c) After receiving the decision to cancel classification result, the authority that has received the application for declaration of applied standards or application for registration number (hereinafter referred to as “application for registration number”) shall review the registration numbers issued, and follow procedures for revoking the registration numbers of medical devices[70] that have been issued using classification results specified in the Ministry of Health’s decision to cancel classification results of such medical devices[71].

Article 7. Handling of medical devices[72] whose classification results are cancelled

1. In case a medical device[73] whose classification result is cancelled has not been granted the registration number:

a) The applicant for registration number shall send a written request to the receiving authority to suspend procedures for issuance of registration number.

b) After receiving the written request mentioned in Point a of this Clause or the decision to cancel medical device[74] classification result, the receiving authority shall refuse to issue the registration number.

2. In case a medical device[75] whose classification result is cancelled has been granted the registration number but has not been granted customs clearance:

a) The registration number holder shall stop following procedures for customs clearance, request the checkpoint customs authority to suspend procedures for customs clearance, and request the registration number issuer to revoke the issued registration number.

b) After receiving the written request from the registration number holder or the decision to cancel medical device[76] classification result, the customs authority shall suspend procedures for customs clearance, and the registration number issuer shall follow procedures to revoke the issued registration number.

3. In case a medical device[77] whose classification result is cancelled has been granted the registration number and customs clearance but has not been purchased:

a) The registration number holder shall:

- Stop placing the medical device[78] granted the registration number using the classification result that has been cancelled on the market, and adopt measures to recall it[79];

- Send a report to the customs authority that granted customs clearance in which the quantity of medical device[80] granted customs clearance must be specified, and do not follow procedures for import of the following shipments until the relevant medical device[81] is granted a new registration number using the correct classification result;

- Send a report to the registration number issuer specifying the quantity of medical device[82] granted customs clearance and the sale contracts (if any);

- Reapply for the registration number.

b) After receiving the registration number holder’s written request or the decision to cancel classification result of the medical device[83]:

- The customs authority shall refuse to grant customs clearance;

- The registration number issuer shall revoke the issued registration number.

4. If the medical device[84] has been sold to health facilities:

a) The registration number holder shall:

- Send a report to the registration number issuer specifying the quantity of medical device[85] sold to health facilities;

- Send written notices to health facilities that purchased the medical device[86].

b) If the medical device[87] granted the registration number using incorrect classification result does not pose threats to patients’ heath, health facilities may keep using the medical device[88] and the registration number holder shall supplement the documents about registration of the medical device[89] after a new registration number is granted.

c) If the medical device[90] granted the registration number using incorrect classification result is found to pose threats to patients’ heath, health facilities shall stop using the medical device[91] and the registration number holder shall implement necessary measures for ensuring normal operation of such health facilities.

Chapter III

MANUFACTURE OF MEDICAL DEVICES[92]

Article 8. Requirements for quality control by medical device[93] manufacturer

1. The manufacturer’s quality control system must comply with ISO 13485.

2. A manufacturer of the medical device[94] that contains narcotic substances and precursors shall, in addition to the requirements laid down in Clause 1 of this Article, also meet the following requirements:

a) It has a system for monitoring the dispatching, warehousing, inventory and use of narcotic substances and precursors, the medical devices[95], and raw materials containing narcotic substances and precursors;

b) Medical devices[96] and raw materials containing narcotic substances and precursors must be safely stored in a separate area or warehouse.

Article 9. Application for declaration of eligibility for manufacture of medical devices[97]

An application for declaration of eligibility for manufacture of medical devices[98] includes:

1. The declaration of eligibility for manufacture of medical devices.

2. Certificate of conformity with ISO 13485 quality control standards issued by the conformity assessment body in accordance with regulations of law.

3. The documents proving the satisfaction of eligibility requirements specified in Clause 2 Article 8 of this Decree.

Article 10. Requirements for application for declaration of eligibility for manufacture of medical devices[99]

1. A set of application for declaration of eligibility for manufacture of medical devices[100] shall be prepared. To be specific:

a) Documents included in the application must be clearly printed and arranged in the order prescribed in Article 9 of this Decree; space must be provided between content parts in a document; the application must have covers and a list of documents.

b) Documents in a language other than English and Vietnamese must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

2. Requirements for some documents in an application for declaration of eligibility for manufacture of medical devices:

a) The original or certified true copy or copy bearing the applicant’s certification of the Certificate of conformity with quality control standards.

b) The originals bearing the applicant’s certification of documents proving its satisfaction of eligibility requirements as set out in Clause 2 Article 8 of this Decree.

Article 11. Procedures for declaration of eligibility for manufacture of medical devices[101]

1. Before initiating manufacture of medical devices[102], the medical device [103] manufacturer shall submit an application for declaration of eligibility for manufacture of medical devices to the Department of Health of province where its factory is located (the manufacturer’s factory is specified in the certificate of quality control system). Where there are multiple factories located in different provinces, application for each province shall be submitted.

2. After receiving the application (including application fee receipt as prescribed by the Ministry of Finance), the Department of Health of province where the medical device[104] manufacturing factory is located shall publish on the Portal on management of medical devices[105] all information about and application for declaration of eligibility for manufacture of medical devices[106].

3. During its operation, the manufacturer shall prepare a notice of changes which is accompanied by supporting documents for such changes, and update such documents to its application for declaration of eligibility for manufacture of medical devices published on Portal on management of medical devices[107] within 03 working days from the occurrence of such changes.

Chapter IV

CLINICAL TRIALS FOR MEDICAL DEVICES[108]

Article 12. Clinical trial phases for medical devices[109]

1. A clinical trial for medical devices other than in vitro diagnostic (IVD) medical devices has the following phases:

a) Phase 1: preliminary study on the safety of medical device for patients and its usability for medical practitioners and healthcare workers.

b) Phase 2: confirmatory study to ascertain and establish evidence for the safety and efficacy of medical device;

c) Phase 3: post-market study which is conducted after the medical device has been placed on market to collect additional evidences on the safety and efficacy of the medical device when used at the request of competent authorities.

2. A clinical trial for IVD medical devices has the following phases:

a) Phase 1: evaluation of clinical performance of the IVD medical device;

b) Phase 2: post-market study which is conducted after the IVD medical device has been placed on market to collect additional evidences on the safety and efficacy of the IVD medical device when used at the request of a competent authority.

Article 13. Exemption from clinical trials or some clinical trial phases before applying for registration number for medical devices[110]

1. The following medical devices will be exempt from clinical trials:

a) Class-A or Class-B medical devices;

b) Class-C or Class-D medical devices for which registration number has been granted by the Ministry of Health of Vietnam or competent authorities of foreign countries.

2. Class-C or Class-D medical devices which are domestically manufactured and are not IVD medical devices will be exempt from the clinical trial Phase 1 if the following requirements are met:

a) The medical device has been manufactured employing a transferred technology which has been used in the manufacturing of a medical device granted registration number by a competent authority, unless such registration number has been revoked;

b) The subject medical device has the same indications for use as the one which has been granted registration number by a competent authority as prescribed in point a of this clause.

3. The exemption from clinical trials or one or some clinical trial phases granted to medical devices used to serve urgent needs for national defense and security, epidemic prevention and control or disaster recovery shall be subject to the Ministry of Health of Vietnam’s decision.

Article 14. Requirements for medical devices undergoing clinical trial [111]

1. All legal and technical requirements set out in documentation requirements for registration of a medical device are met.

2. The investigational medical device must be labeled as “for clinical trial only - not for any other purposes”. The labeling of medical devices shall comply with regulations of law on labeling of goods.

Article 15. Eligibility requirements for institutions conducting medical device clinical trials (hereinafter referred to as “clinical trial institution”) [112]

1. It must meet good clinical practices (GCP) for medical devices in accordance with regulations of the Minister of Health of Vietnam.

2. It is capable of performing scientific research function and has its scope of operation conformable with the medical device subject to the clinical trial.

3. It must be a juridical person whose organizational structure, personnel and finance are independent from the organization or individual whose medical device is subject to the clinical trial.

Article 16. Clinical trial dossier for medical devices [113]

1. A clinical trial dossier for medical device includes: an application for approval of clinical trial, application for approval of changes in clinical trial, and application for approval of clinical trial results. Components of each application will comply with the provisions of Article 17 of this Decree.

2. Requirements for documents:

a) Documents included in the clinical trial dossier for a medical device must be made in Vietnamese. Documents in a language other than Vietnamese must be accompanied by their certified Vietnamese translations.

b) Documents issued by foreign regulatory authorities must bear consular legalization made in accordance with regulations of law on consular legalization, unless such consular legalization procedures are exempted in accordance with regulations of law.

Article 17. Components of a clinical trial dossier for medical device [114]

1. An application for approval of clinical trial includes:

a) An application form made by the organization or individual whose medical device is subject to the clinical trial.

b) An application form made by the clinical trial institution.

c) A competent authority’s written request in respect of the medical device which must undergo the clinical trial phase 3 or the IVD medical device which must undergo the clinical trial phase 2.

d) Documents on the medical device subject to the clinical trial, including:

- Description of the product (general information about the medical device subject to the clinical trial, including: name, characteristics, uses, indications, contraindications, technical specifications and other relevant information);

- Documents on preclinical trial on the investigational medical device, including: study report on the safety and efficacy of the medical device, recommendations for use and storage of the medical device, and study reports on the stability and analytical performance (of IVD medical device) which have been appraised and certified to be satisfactory by the Expert Council of the clinical trial institution or the institution conducting preclinical trial vested with certification task by the Ministry of Health of Vietnam;

- Documents on the clinical trial phases before the requested phase.

dd) Legal documents, including:

- Information on the owner and manufacturer of the medical device, manufacturing process;

- Technical file of the medical device;

- Technical standards applicable to the medical device as announced by its owner;

- Assessment report certifying that the medical device is conformable with technical standards provided by a qualified conformity assessment body as prescribed by law or quality assessment certificate issued by a competent authority of Vietnam;

- The approved user manual, if the medical device undergoes the clinical trial phase 3;

- Certificates of consent given by institutions participating in the trial;

- The clinical trial agreement signed by and between the authority, organization or individual whose medical device is subject to the clinical trial and the clinical trial institution; the research support agreement signed by and between the organization or individual whose medical device is subject to the clinical trial and the sponsor (if any).

e) Scientific documents, including:

- Description of the medical device clinical trial;

- Case Report Form (CRF);

- Informed consent forms given by the clinical trial participants.

g) Administrative documents:

- Academic resume and copy of certificate of completion of a training course in GCP, issued by the Ministry of Health of Vietnam or a qualified GCP training institution, of the principal investigator;

- A record of scientific and ethical assessment prepared by the internal Research Ethics Committee;

- Label of the investigational medical device as prescribed in clause 2 Article 14 of this Decree and its user manual.

2. An application for approval of changes in clinical trial includes:

a) For changes in administrative documents or contact information of the clinical trial institution: this institution shall take charge of reporting such changes to the Ministry of Health of Vietnam for information consolidation.

b) For replacement of the principal investigator or the clinical trial institution, the application shall include:

- An application form for approval of changes in clinical trial;

- Documents proving corresponding changes.

c) For changes other than those specified in points a and b of this Clause, the application shall include:

- An application form for approval of changes in clinical trial;

- Updated versions of the corresponding documents specified in clause 1 of this Article, as changed.

3. An application for approval of clinical trial results includes:

a) An application form for approval of clinical trial results;

b) The copy of the approved clinical trial protocol;

c) A copy of the decision to approve the clinical trial;

d) The record of appraisal of clinical trial results made by the internal Research Ethics Committee;

dd) Report on clinical trial results.

Article 18. Procedures for approving clinical trial [115]

1. The clinical trial institution shall send an application for approval of clinical trial directly, online or by post to the Ministry of Health of Vietnam.

2. The Ministry of Health shall check the validity of the application within 05 working days from the date of receipt. If the application is valid, within 02 working days, the Ministry of Health of Vietnam shall transfer the application to the National Research Ethics Committee for appraisal and ethical considerations. If the application is invalid, within 02 working days, a written notice providing specific instructions for application modification shall be given to the applicant.

The clinical trial institution shall complete the application within 60 days from the receipt of the notice. After this period, procedures for approving clinical trial must be restarted.

3. The National Research Ethics Committee shall conduct review and ethical consideration of the trial on a periodical basis of every month.

4. Within 05 working days from the receipt of the certificate of approval of clinical trial given by the National Research Ethics Committee and the accompanied application, the Ministry of Health of Vietnam shall issue a decision to approve clinical trial.

Article 19. Procedures for approving changes in clinical trial [116]

1. The clinical trial institution shall send an application for approval of changes in clinical trial directly, online or by post to the Ministry of Health of Vietnam.

2. The Ministry of Health shall check the validity of the application within 05 working days from the date of receipt. If the application is invalid, within 02 working days, the Ministry of Health of Vietnam shall give a written notice providing specific instructions for modification until the application is considered valid.

The clinical trial institution shall complete the application within 60 days from the receipt of the notice. After this period, procedures for approving changes in clinical trial must be restarted.

3. For changes specified in point b clause 2 Article 17 of this Decree: within 07 working days from the receipt of adequate and valid application, the Ministry of Health of Vietnam shall issue a decision to approve changes in clinical trial.

4. For changes specified in point c clause 2 Article 17 of this Decree:

a) If the application is valid, within 02 working days, the Ministry of Health of Vietnam shall transfer it to the National Research Ethics Committee for review;

b) The National Research Ethics Committee shall conduct review and ethical consideration of the trial on a periodical basis of every month;

c) Within 05 working days from the receipt of the certificate of approval of changes in clinical trial given by the National Research Ethics Committee and the accompanied application, the Ministry of Health of Vietnam shall issue a decision to approve changes in clinical trial.

Article 20. Procedures for approving clinical trial results [117]

1. The clinical trial institution shall send an application for approval of clinical trial results, which must be made in Vietnamese, directly, online or by post to the Ministry of Health of Vietnam.

2. The Ministry of Health shall check the validity of the application within 05 working days from the date of receipt. If the application is valid, within 02 working days, the Ministry of Health of Vietnam shall transfer it to the National Research Ethics Committee for review. If the application is invalid, within 02 working days, the Ministry of Health of Vietnam shall give a written notice providing specific instructions for modification to the clinical institution.

The clinical trial institution shall complete the application within 60 days from the receipt of the notice. After this period, procedures for approving clinical trial results must be restarted.

3. The National Research Ethics Committee shall conduct review and ethical consideration of the trial on a periodical basis of every month.

4. Within 05 working days from the receipt of the certificate of approval of clinical trial results given by the National Research Ethics Committee and the accompanied application, the Ministry of Health of Vietnam shall issue a decision to approve clinical trial results.

Chapter V

PLACEMENT OF MEDICAL DEVICES ON MARKET [118]

Section 1. REGISTRATION NUMBER, CONDITIONS FOR PLACEMENT ON MARKET AND REQUIREMENTS FOR ORGANIZATIONS DECLARING APPLIED STANDARDS OR APPLYING FOR REGISTRATION NUMBER

Article 21. Registration number of medical devices [119]

1. Registration number of a medical device[120] is:

a) The number of declaration of applied standards of Class-A or Class-B medical device[121];

b) The number of certificate of registration of Class-C or Class-D medical device[122].

2. [123] Registration number holder is the organization that declares applied standards for medical devices[124] or is issued with the certificate of registration of medical devices[125] in accordance with the provisions of this Decree.

3. Validity of registration number: The medical device[126] registration number shall be valid for indefinite term, except cases where a medical device[127] registration number is issued according to regulations on issuance of emergency use registration number for medical devices[128] to serve epidemic prevention and control, and disaster recovery purposes. Based on the application for issuance of emergency use registration number for the medical device[129], the Minister of Health shall decide specific period validity of the issued registration number.

Article 22. Conditions for placement of medical devices on market [130] [131]

1. A medical device[132] may be placed on the market if it meets the following conditions:

a) It has been granted registration number, marketing authorization number, certificate of registration or import license in accordance with regulations on management of medical devices[133] or in the cases specified in Point d Clause 2 Article 76 of this Decree, except the following cases:

- The medical device is liquidated as prescribed by laws;

- The medical device reaches its expiry date;

- The defect that is harmful to users’ health cannot be repaired as prescribed in Clause 4 Article 34 of this Decree;

- The use of medical device is prohibited by a regulatory authority.

b) Its label contains adequate information in accordance with regulations of law on labelling of goods;

c) Instructions for use of the medical device[134] are given in Vietnamese language;

d) Information about warranty center, conditions and time for warranty, except disposable medical devices[135] defined by product owners[136] or cases where there are documents proving that the medical device is not under warranty.

2. If the import license is available as prescribed in Points a, b, c, d and dd Clause 1 Article 48 of this Decree, the satisfaction of the condition in Point d Clause 1 of this Article is not required.

3. If the information specified in Point c and Point d Clause 1 of this Article is not provided upon the medical device[137] itself, it must be provided in the form of electronic information for which instructions for search must be available on the label of the medical device[138].

Article 23. Requirements for declaration of applied standards or issuance of certificate of registration for medical devices [139]

1. Requirements for declaration of applied standards or issuance of certificate of registration for medical devices[140]:

a) The medical device is manufactured domestically by a manufacturer that has made declaration of eligibility for manufacture of medical devices[141];

b) If the medical device[142] is imported, it has been manufactured by a manufacturer that is granted certificate of conformity with ISO 13485 quality control standards, and permitted to be sold in every country in the world;

c) The medical device comply with national technical regulations or standards declared by the manufacturer.

2. Applications for declaration of applied standards or registration of medical devices[143] shall be rejected in the following cases:

a) The medical device [144] falls in the case specified in Clause 1 Article 37 of this Decree;

b) The medical device[145] is recalled under Clauses 1, 3 Article 38 of this Decree.

3. If a medical device[146] registration number is revoked under Clause 2 Article 38 of this Decree, an application for declaration of applied standards or registration of that medical device submitted within 12 months from the date of the revocation decision shall be rejected.

Article 24. Cases of exemption from requirements for declaration of applied standards and registration

1. The medical device[147] is used for research, testing, inspection, experiment, performance evaluation, instruction for use or modification of medical devices[148] only.

2. The medical device[149] is imported into Vietnam to serve the purposes of emergency aid or provision of humanitarian medical services or display at trade fair, exhibition or product launch event or use as gift given to a health facility or for medical treatment of a particular individual or to serve special diagnosis of a health facility.

3. Unregistered medical devices[150] are imported to serve epidemic prevention and control, disaster recovery and cannot be replaced by any other medical devices[151] sold on the market.

4. Medical devices[152] are manufactured in Vietnam to serve the purposes of export or display at an overseas trade fair or exhibition only.

Article 25. Requirements for organizations declaring applied standards or applying for registration number for medical devices [153]

1. Organizations that declare applied standards or apply for registration number for medical devices [154] include:

a) Vietnamese enterprises, cooperatives or household businesses that are product owners [155];

b) Vietnamese enterprises, cooperatives or household businesses that are authorized by product owners[156];

c) Permanent representative offices in Vietnam of foreign traders that are product owners[157] or authorized by product owners[158].

2. The organization declaring applied standards or applying for registration of a medical device[159] must establish and maintain a warranty center in Vietnam or sign a contract with a qualified organization to provide warranty for medical devices[160], except disposable medical devices[161]  as defined by product owners [162] or cases where there are documents proving that the medical device is not under warranty.

Where the organization declaring applied standards or applying for registration of a medical device[163] is the entity prescribed in Point c Clause 1 of this Article, the product owner[164] must establish and maintain a warranty center in Vietnam or sign a contract with a qualified organization to provide warranty for medical devices[165], except disposable medical devices[166] as defined by product owners[167] or cases where there are documents proving that the medical device is not under warranty.

The warranty center must be granted certificate of eligibility to provide warranty by the product owner[168].

Section 2. DECLARATION OF APPLIED STANDARDS FOR CLASS-A OR CLASS-B MEDICAL DEVICES [169]

Article 26. Application for declaration of applied standards

An application for declaration of applied standards for Class-A or Class-B medical device[170] consists of:

1. The declaration of applied standards for Class-A or Class-B medical device[171].

2. An unexpired certificate of conformity with ISO 13485 quality control standards.

3. The unexpired letter of authorization given by the product owner[172] to the organization that applies for declaration of applied standards, except the case specified in Point a Clause 1 Article 25 of this Decree.

4. The certificate of eligibility to provide warranty services granted by the product owner[173], except disposable medical devices[174] defined by product owners[175] or cases where there are documents proving that the medical device is not under warranty.

5. A synopsis of technical description of the medical device[176]  in Vietnamese, accompanied by technical documents describing functions and specifications of the medical device[177] issued by the product owner[178].

With regard to in-vitro reagents, calibrators and control materials, the synopsis of technical description in Vietnamese must be accompanied by documents on materials and safety of the product, manufacturing process, pre-clinical and clinical trial reports including stability report.

6. The certificate of conformity or product standard sheet provided by the product owner[179].

With regard to domestically manufactured medical devices[180], the assessment record of chemical, physical and microbiological indicators and other indicators provided by a qualified conformity assessment body in accordance with regulations of law on conformity assessment is required; if it is an IVD medical device[181], certificate of quality assessment issued by a competent authority of Vietnam is required. The assessment result must be conformable with the standards declared by the product owner[182].

7. User manual of the medical device[183].

8. Sample of the label for the medical device[184] sold in Vietnam.

9. An unexpired CFS (for imported medical devices[185]).

Article 27. Requirements for application for declaration of applied standards

1. 01 set of application for declaration of applied standards shall be prepared.

2. Requirements for some documents included in the application:

a) Certificate of conformity with quality control standards: original copy or certified true copy or copy bearing certification of the declarant.

If the Certificate of conformity with quality control standards is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

b) Product owner’s[186] letter of authorization and certificate of eligibility to provide warranty:

- For a domestically manufactured medical device[187]: original copy or certified true copy;

- For an imported medical device[188]: Either the document bearing consular legalization or the certified true copy thereof is accepted.

c) Certificate of conformity or product standard sheet provided by the product owner[189]: original copy or copy bearing certification of the declarant.

If the product standard sheet is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

d) User manual of the medical device[190]: Vietnamese copy bearing certification of the declarant, accompanied by English version issued by the product owner[191], in case of imported medical devices[192]. If the user manual is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

dd) Sample label: the sample label bearing certification of the declarant. The sample label must meet requirements laid down in regulations of law on labeling of goods.

e) CFS: Either the document bearing consular legalization or the certified true copy thereof is accepted.

If the CFS is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

g) Record of chemical, physical and microbiological indicators and other indicators provided by a qualified conformity assessment body in accordance with regulations of law on conformity assessment or certificate of quality assessment issued by regulatory authority of Vietnam (for an IVD medical device[193]):  original copy or certified true copy.

Article 28. Procedures for declaration of applied standards

1. Before placing Class-A or Class-B medical device[194] on the market, the organization declaring applied standards (declarant) shall submit an application for declaration to the Department of Health of province where it is headquartered.

2. After receiving the application (including application fee receipt as prescribed by the Ministry of Finance), the Department of Health of province where the declarant of the medical device[195] is headquartered shall publish the number of declaration of applied standards for Class-A or Class-B medical device[196] on the Portal on management of medical devices[197] and applications for declaration of applied standards, except the documents specified in Clause 5 Article 26 of this Decree.

3. In case of changes in the product owner[198], class of the medical device[199], category, intended purposes or indications for use of the medical device, or provision of additional information about the manufacturer or product code, the registration number holder shall reapply for declaration of applied standards in accordance with regulations of this Decree.

4. During the placement of the medical device [200] on the market, the product owner that has declared applied standards for the medical device shall prepare written notice of changes, accompanied by supporting documents, and update them to its application for declaration of applied standards published on the Portal on management of medical devices[201] within 05 working days from the occurrence of any of the following changes:

a) Change in address of the product owner[202] or registration number holder of the medical device[203];

b) Change in name of the registration number holder or product owner[204];

c) Change in the medical device[205] manufacturer’s name or address;

d) Change in packaging specifications;

dd) Change in the warranty center;

e) Change in the label or user manual without changing intended purposes or indications for use. Where the medical device[206] has been manufactured before the product owner makes a notice of change in the label, the device may be placed on the market with the information declared at the date of manufacture;

g) Reduction of factories, category or product code.

Section 3. REGISTRATION OF CLASS-C OR CLASS-D MEDICAL DEVICES [207]

Article 29. Methods of registration

1. A new registration number shall be issued to a medical device[208] in the following cases:

a) The medical device[209] does not have any registration number;

b) The medical device[210] has been granted registration number but has any of the following changes:

- A change in the product owner[211]; class of the medical device[212]; category, intended purposes or indications for use; quality standards; addition of factory or product code; raw materials that affect the functions of IVD medical device[213] or disposable medical device[214]; concentration, content, composition of raw materials that are active ingredients incorporated in the medical device[215] to support in medical treatment;

- A change other than those specified in Clause 7 Article 32 of this Decree.

2. Cases of quick issuance of new registration number to a medical device[216]:

a) The medical device has been granted the CFS (Certificate of Free Sale) or marketing authorization by one of the following authorities or countries: U.S. Food and Drug Administration (FDA), Australian Therapeutic Goods Administration (TGA), Health Canada, Japanese Ministry of Health, Labour and Welfare (MHLW) or Pharmaceuticals and Medical Devices Agency, EU state members, UK, Switzerland, National Medical Products Administration (NMPA) - China, Ministry of Food & Drug Safety (MFDS) - Korea, or other CFS issuing authorities recognized by Vietnam (hereinafter referred to as “reference countries”);

b) The medical device has been granted the import license or registration number or certificate of registration in the commercial form in Vietnam, unless revoked before the effective date of this Decree;

3. Issuance of emergency use registration number for medical devices[217] to serve epidemic prevention and control, and disaster recovery purposes in emergency cases included in the relevant List issued by the Minister of Health and in one of the following cases:

a) The medical device has been placed on the market or granted registration number for emergency use in one of the reference countries;

b) The medical device is included in the list of products for emergency use announced by the World Health Organization (WHO);

c) The medical device is included in the list of common products announced by the EU Health Security Committee (EUHSC);

d) The medical device has been granted the registration number or import license in the commercial form in Vietnam, unless revoked before the effective date of this Decree;

dd) The medical device[218] is domestically manufactured in the form of technology transfer in one of the cases specified in Points a, b, c or d of this Clause;

e) The medical device[219] is domestically manufactured in the form of processing in one of the cases specified in Points a, b, c or d of this Clause.

Article 30. Application for issuance of a new registration number

1. An application for issuance of a new registration number of a medical device[220] without a national technical regulation consists of:

a) The application form for issuance of a new registration number.

b) An unexpired certificate of conformity with ISO 13845 quality control standards.

c) The unexpired letter of authorization given by the product owner[221] to the organization that applies for registration, except the case specified in Point a Clause 1 Article 25 of this Decree.

The certificate of eligibility to provide warranty services granted by the product owner[222], except disposable medical devices[223] defined by product owners[224] or cases where there are documents proving that the medical device is not under warranty.

dd) The unexpired CFS (for imported medical devices[225]).

e) ASEAN Common Submission Dossier Template on the medical device[226] (hereinafter referred to as “CSDT”).

g) Certificate of conformity.

h)[227] Certificate of results of clinical trial for medical device issued by a competent authority, if the medical device is subject to a clinical trial as prescribed in clause 2 Article 94 of the Law on Medical Examination and Treatment.

2. An application for issuance of a new registration number for the medical device[228] that is a measuring device whose sample requires approval as prescribed by the law on measurement consists of:

a) The application form for issuance of a new registration number.

b) Decision on approval of the sample medical device.

c) [229] The documents specified in Points b, c, d, dd, e and h clause 1 of this Article.

3. An application for quick issuance of a new registration number for the medical device[230] as prescribed in Clause 2 Article 29 of this Decree consists of:

a) The application form for issuance of a new registration number.

b) The documents specified in Points b, c, and d Clause 1 of this Article.

c) The CFS issued by one of the reference countries in the case prescribed in Point a Clause 2 Article 29 of this Decree.

d) The CFS of the imported medical device[231] and import license or registration number or certificate of registration in the case prescribed in Point b Clause 2 Article 29 of this Decree.

dd) Certificate of quality assessment issued by a competent authority of Vietnam for IVD medical device[232], except the following cases:

- The medical device is included in List A or List B of Annex 2 of the European Union Regulation of In Vitro Diagnostic Medical Devices[233] and has been granted CFS by one of EU Member States, UK or Switzerland;

- The medical device is included in List A or List B of Annex 2 of the European Union Regulation of In Vitro Diagnostic Medical Devices[234] and has been granted the Marketing Authorization by one of the reference countries;

- The medical device is not included in List A or List B of Annex 2 of the European Union Regulation of In Vitro Diagnostic Medical Devices[235] but has been granted the Marketing Authorization by one of the reference countries;

- The medical device is included in the list of medical devices announced by the Minister of Health.

e) The CSDT.

4. An application for issuance of a new registration number for the medical device[236] in an emergency case prescribed in Clause 3 Article 29 of this Decree consists of:

a) The application form for issuance of a new registration number.

b) The documents specified in Points b, c, and d Clause 1 of this Article.

c) The CFS or license for emergency use of the imported medical device[237].

d) The contract for technology transfer in the case prescribed in Point dd Clause 3 Article 29 of this Decree.

dd) The processing contract in the case prescribed in Point e Clause 3 Article 29 of this Decree.

e) Certificate of inspection or certificate of quality assessment issued by one of the inspection/assessment bodies in the list published on the Ministry of Health’s website if the medical device falls in one of the following cases:

- The medical device[238] is domestically manufactured;

- The medical device[239] has been granted permit for placement on the market or emergency use by a competent authority of one of EU Member States, UK or Switzerland but is not included in the EUHSC’s list of common products.

g) The CSDT.

h) [240] Certificate of results of clinical trial for medical device issued by a competent authority, if the medical device is subject to a clinical trial as prescribed by law.

5. An application for issuance of a new registration number for another medical device[241] consists of:

a) The application form for issuance of a new registration number.

b) The documents specified in Points b, c, d and dd Clause 1 of this Article.

c) The CSDT inspection record given by the authority designated by the Minister of Health, accompanied by the CSDT.

d) With regard to IVD medical device[242] that is reagent, calibrator or control material: certificate of quality issued by a competent authority of Vietnam.

dd) With regard to chemicals or preparations only used for disinfection of medical devices[243]: Report on testing for active ingredients for disinfection given by the entity that has declared its eligibility to provide testing services in accordance with regulations of law on management of insecticidal and germicidal chemicals and preparations for medical and household use; report on trials to evaluate biological efficacy and side effects of product on trial participants given by the entity that has declared its eligibility to conduct trial in accordance with regulations of law on management of insecticidal and germicidal chemicals and preparations for medical and household use.

e) [244] Certificate of results of clinical trial for medical device issued by a competent authority, if the medical device is subject to a clinical trial as prescribed in clause 2 Article 94 of the Law on Medical Examination and Treatment.

Article 31. Requirements for application for issuance of a new registration number

1. Requirements for some documents included in the application for registration:

a) Certificate of conformity with quality control standards: original copy or certified true copy or copy bearing the applicant’s certification.

If the Certificate of conformity with quality control standards is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

b) Product owner’s[245] letter of authorization and certificate of eligibility to provide warranty:

- For a domestically manufactured medical device[246]: original copy or certified true copy;

- For an imported medical device[247]: Either the document bearing consular legalization or the certified true copy thereof is accepted.

c) CFS: Either the document bearing consular legalization or the certified true copy thereof is accepted.

If the CFS is made neither in English nor in Vietnamese, it shall be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

d) Certificate of quality assessment, testing report, trial report and CSDT inspection record: original copy or certified true copy or copy bearing the applicant’s certification.

dd) CSDT: copy bearing the applicant’s certification. If the CSDT is not written in English or Vietnamese, it must be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

2. Requirements for some documents included in an application for registration of medical device for emergency use as prescribed in Clause 3 Article 29 of this Decree:

Documents issued by foreign competent authorities must bear consular legalization. Where a document does not bear consular legalization:

a) With regard to the letter of authorization: its original copy accompanied by certification shall be submitted.

b) With regard to the document specified in Point c Clause 4 Article 30 of this Decree: it must indicate the link for searching information on placement on the market or license for use of the medical device[248] on the licensing authority’s website and be accompanied by a document providing the link for searching information of the applicant. Information on placement of the device on the market found on the website must include at least the following information in English: name, category, manufacturer and manufacturing country.

Article 32. Receipt and processing of application for registration of medical devices [249]

1. The applicant shall submit an application for registration number to the Ministry of Health through the Portal on management of medical devices[250].

2. With regard to an application for registration number for the medical device[251] in the case prescribed in Clause 1 or Clause 2 Article 30 of this Decree:

a) If the application is satisfactory, the Minister of Health shall process it within 30 days from the receipt of the adequate and valid application (including the application fee receipt as prescribed by the Ministry of Finance). If an application is refused, a written response indicating reasons for such refusal shall be provided.

b) If the application is not satisfactory, the Ministry of Health shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant within 25 days from the receipt of the application.

c) The applicant shall comply with the request for modification and send modified documents to the Ministry of Health as requested.

If the modified application is still unsatisfactory, the Ministry of Health shall continue sending a request for modification to the applicant as prescribed in Point b Clause 2 of this Article.

If the applicant fails to provide the modified application within 90 days from the receipt of the Ministry of Health’s request for modification or the application is still unsatisfactory after 05 modification times, the application shall be rejected.

3. With regard to an application for registration number for the medical device[252] in the case prescribed in Clause 3 Article 30 of this Decree:

a) If the application is satisfactory, the Minister of Health shall process it within 15 days from the receipt of the adequate and valid application (including the application fee receipt as prescribed by the Ministry of Finance). If an application is refused, a written response indicating reasons for such refusal shall be provided.

The Ministry of Health shall only appraise the legal documents specified in Points b, c, d and dd Clause 1 Article 30 of this Decree. The applicant shall assume responsibility for the accuracy and legitimacy of other documents included in the application which may be inspected by the Ministry of Health after issuing the registration number.

b) If the application for registration number is not satisfactory, the Ministry of Health shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant within 12 days from the receipt of the application.

c) [253] The applicant shall comply with the request for modification and send modified documents to the Ministry of Health as requested.

If the modified application is still unsatisfactory, the Ministry of Health shall continue sending a request for modification to the applicant as prescribed in Point b Clause 3 of this Article.

If the applicant fails to provide the modified application within 90 days from the receipt of the Ministry of Health’s request for modification or the application is still unsatisfactory after 03 modification times, the application shall be rejected.

4. With regard to an application for registration number for the medical device [254] in the case prescribed in Clause 4 Article 30 of this Decree:

a) If the application is satisfactory, the Minister of Health shall process it within 10 days from the receipt of the adequate and valid application (including the application fee receipt as prescribed by the Ministry of Finance). If an application is refused, a written response indicating reasons for such refusal shall be provided.

The Ministry of Health shall only appraise the legal documents specified in Points b, c, d, dd and e Clause 4 Article 30 of this Decree. The applicant shall assume responsibility for the accuracy and legitimacy of other documents included in the application which may be inspected by the Ministry of Health after issuing the registration number.

b) If the application for registration number is not satisfactory, the Ministry of Health shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant within 08 days from the receipt of the application.

c) The applicant shall comply with the request for modification and send modified documents to the Ministry of Health as requested.

If the modified application is still unsatisfactory, the Ministry of Health shall continue sending another request for modification to the applicant as prescribed in Point b Clause 4 of this Article.

If the applicant fails to provide the modified application within 90 days from the receipt of the Ministry of Health’s request for modification, the application shall be rejected.

5. With regard to an application for registration number for the medical device[255] in the case prescribed in Clause 5 Article 30 of this Decree:

a) If the application is satisfactory, the Minister of Health shall process it within 45 days from the receipt of the adequate and valid application (including the application fee receipt as prescribed by the Ministry of Finance). If an application is refused, a written response indicating reasons for such refusal shall be provided.

b) If the application for registration number is not satisfactory, the Ministry of Health shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant within 40 days from the receipt of the application.

c) The applicant shall comply with the request for modification and send modified documents to the Ministry of Health as requested.

If the modified application is still unsatisfactory, the Ministry of Health shall continue sending another request for modification to the applicant as prescribed in Point b Clause 5 of this Article.

If the applicant fails to provide the modified application within 90 days from the receipt of the Ministry of Health’s request for modification or the application is still unsatisfactory after 05 modification times, the application shall be rejected.

6. Within 01 working day from the issuance of the registration number, the Ministry of Health shall publish the following information on the Portal on management of medical devices[256], including:

a) Name, category, manufacturer and manufacturing country of the medical device[257];

b) Registration number of the medical device [258];

c) Name and address of the product owner[259];

d) Name and address of the registration number holder;

dd) Name and address of the warranty center[260];

e) The documents included in the application for registration of the medical device[261], except the documents specified in Point e Clause 1 and Point c Clause 5 of Article 30;

g) Intended purposes of the medical device[262].

7. During the placement of the medical device[263] on the market, the registration number holder shall prepare written notice of changes, accompanied by supporting documents, and update them to its application for registration number published on the Portal on management of medical devices[264] within 10 working days from the occurrence of any of the following changes:

a) Change in address of the product owner[265] or registration number holder of the medical device[266];

b) Change in name of the registration number holder or product owner[267];

c) Change in the medical device[268] manufacturer’s name or address;

d) Change in packaging specifications;

dd) Change in the warranty center;

e) Change in the label or user manual without changing intended purposes, indications for use, functions or performance of the medical device[269]. Where the medical device[270] has been manufactured before the registration number holder makes a notice of change in the label, the device may be placed on the market with information registered and updated at the date of manufacture.

Article 32a. Priority in processing applications for marketing authorization for medical devices [271]

1. Processing of applications for registration number for medical devices shall be prioritized in the following cases:

a) Domestically manufactured medical devices;

b) Imported medical devices for which applications for import license have been submitted to the Ministry of Health of Vietnam before January 01, 2022 but not yet been processed, and for which applications for registration number have also been submitted;

c) Chemicals and preparations which are only used for decontamination of medical devices;

d) Medical devices which are employed for performing new techniques or methods in medical examination and treatment under official approvals given by the Ministry of Health of Vietnam but for which registration numbers have not yet been granted in Vietnam;

dd) Medical devices for which there is no change in information on type and HS code on the unexpired import license or decision to grant registration number or certificate of registration but which is subject to one of the following cases:

- Changes in the importer or the applicant for registration number;

- Changes in name of the manufacturer but its address remains unchanged;

- Changes in name of the medical device’s owner but its address remains unchanged;

e) A medical device for which there is no information on its type on the unexpired import license or decision to grant registration number or certificate of registration but information on product code on the same document is changed due to terminated manufacturing of products of the previous code;

g) The Ministry of Health of Vietnam has given a document rejecting, or requesting modification of, the submitted application for registration number due to validity of one of the following documents: letter of authorization of the registration number holder, certificate of conformity with ISO 13485 quality standards, or certificate of registration for the medical device;

h) The Ministry of Health of Vietnam has given a document rejecting, or requesting modification of, the submitted application for registration number due to the name of the medical device.

2. Procedures for priority processing of an application for registration number:

a) The applicant shall submit the following documents to the Ministry of Health of Vietnam:

- An application form for priority processing of the application which must indicate adequate information on the code of the submitted application for registration number on the website of the Ministry of Health of Vietnam;

- Corresponding support documents for the cases specified in clause 1 of this Article;

b) The Ministry of Health of Vietnam shall consider permitting the priority processing of application for registration number for medical device within 10 working days from its receipt of the application form for priority processing and accompanied documents. If an application is refused, a written response indicating reasons for such refusal shall be provided.

3. Applications for license to import medical devices prescribed in point b clause 1 of this Article for which certificate of registration has been issued shall not be rejected.

Section 4. POST-MARKET MANAGEMENT OF MEDICAL DEVICES[272] AND HANDLING OF MEDICAL DEVICES [273] IN SOME SPECIFIC CASES

Article 33. Documents for post-market management of medical devices [274]

The registration number holder shall organize and manage the tracing of origin of medical devices[275] placed on the market and fully retain at least the following documents to serve its management of medical devices:

1. The application for issuance of registration number for medical device[276] of which the following documents shall be retained in physical form:

a) The letter of authorization given by the product owner[277] to the organization that applies for registration, except the case specified in Point a Clause 1 Article 25 of this Decree;

b) The certificate of eligibility to provide warranty granted by the product owner[278], except disposable medical devices[279] defined by product owners[280] or cases where there are documents proving that the medical device is not under warranty;

c) The CFS or marketing authorization.

2. Distribution records (if the registration number holder is a representative office, it shall not be required to retain these records but must request the entity that it authorizes to import the medical device to perform this responsibility).

3. Records of adverse events, complaints and corrective actions that indicate name, category, quantity and batch number of the medical device[281], especially those[282] that are defective or unsafe for users.

4. Medical device[283] quality management documents, including:

a) The Certificate of Origin made in accordance with regulations of law on origin of goods;

b) The certificate of quality of each medical device batch issued by the product owner[284] or manufacturer whose name is specified in the application for registration of the medical device[285];

c) Record of inspection of medical device[286] in respect of the medical devices[287] specified in Clause 1 Article 55 of this Decree;

d) Technical documents serving the repair and maintenance of the medical device[288], except disposable medical devices[289] defined by product owners[290] or cases where there are documents proving that the medical device is not under warranty;

dd) Instructions for use of the medical device[291] are given in Vietnamese;

e) Information about warranty center, conditions and time for warranty, except disposable medical devices[292] defined by the product owner[293] or cases where there are documents proving that the medical device is not under warranty.

Article 34. Handling of medical devices [294] that pose a serious threat to public health or may cause death of users

1. In case where a Vietnamese or international competent authority issues a warning against a medical device[295] that poses a serious threat to public health or may cause death of users, the registration number holder shall inform health facilities using such medical device[296] of the warning and carry out investigations within 30 days from the receipt of such warning. If the investigation cannot be finished within 30 days, a report specifying reasons and solutions for ensuring safety of users shall be submitted to the Ministry of Health.

2. In case the medical device[297] specified in Clause 1 of this Article is certified a defective medical device[298] which may affect the health of users, the registration number holder shall:

a) Suspend the placement of the batch of medical devices[299] on the market.

b) Send a written notice to the Ministry of Health, distributors and users of such medical device[300]. The notice shall specify the batch number and the defect, and whether or not such defect can be repaired.

c) Establish a plan for repairing or recalling the batch of defective medical device[301].

d) Send a report to the Ministry of Health after the repair or recall of defective medical devices[302] is completed.

3. If the defect in the medical device[303] can be repaired:

a) Within 03 working days from the receipt of the notice from the registration number holder of the medical device[304], the Ministry of Health shall issue a decision to suspend the placement of the batch of defective medical devices[305] on the market.

A suspension decision shall contain the following information:

- Name of the medical device[306] suspended from placement on the market;

- Batch number of the medical device[307];

- Registration number of the medical device[308].

b) After a decision to suspend the placement of the batch of medical devices[309] on the market is issued, the registration number holder shall repair the defect, which causes adverse effects to the health of users.

c) After completing the repair of the defect, the registration number holder shall send a report accompanied by the inspection record (for the medical device310] specified in Clause 1 Article 55 of this Decree) or commitment to maintain quality of the medical device[311] after the defect has been repaired (for other medical device[312]) to the Ministry of Health.

d) Within 20 working days from the receipt of the report on repair of the defect in the batch of medical devices[313] submitted by the registration number holder, the Ministry of Health shall issue a decision to lift the suspension of such batch of medical devices[314]. If the Ministry of Health refuses to lift the suspension, it shall respond and explain in writing.

4. If the defect in the medical device[315] cannot be repaired:

a) The Ministry of Health shall issue a decision to recall the entire batch of defective medical devices[316].

A recall decision shall contain the following information:

- Name of the medical device[317] to be recalled;

- Batch number of the medical device[318] to be recalled;

- Registration number of the medical device[319].

b) The registration number holder shall recall the entire batch of defective medical devices[320] by the deadline imposed by a competent authority and pay the cost of recall of the entire batch of defective medical devices[321].

c) If the registration number holder fails to complete the recall of the batch of defective medical devices[322] by the deadline imposed by the competent authority, the recall will be enforced in accordance with regulations of law on handling of administrative violations.

Article 35. Actions against adverse events (AE) associated with medical devices [323] that have caused harm to health of users

1. In case an AE associated with the use of a medical device[324] has occurred and poses a serious threat to public health or has caused death of a user, the registration number holder shall:

a) Post a notice of the AE on the registration number holder’s website (if any) and send written notices of the AE to the Ministry of Health, traders and users of that batch of medical devices[325];

b) Suspend the placement on the market of the batch of medical devices[326]  involving the AE;

c) Initiate an investigation into the causes of the AE;

d) Send a report to the Ministry of Health after the investigation result is available. If the AE is caused by a defect in the medical device[327], the report shall describe the defect and whether it can be repaired. Carry out repair of the defect or recall of the batch of defective medical devices[328], submit a report to the Ministry of Health after completing the repair or recall of the batch of defective medical devices[329].

2. In case the AE associated with the use of a medical device[330] does not cause death but has caused serious harm to the health of users, the registration number holder shall:

a) Send a notice of the AE to the Ministry of Health;

b) Initiate an investigation into the causes of the AE;

c) Send a report to the Ministry of Health after the investigation result is available. If the AE is caused by a defect in the medical device[331], the report shall describe the defect and whether it can be repaired. Carry out repair of the defect or recall of the batch of defective medical devices[332], submit a report to the Ministry of Health after completing the repair or recall of the batch of defective medical devices[333].

3. The medical device[334] that has a defect that causes harm to health of users shall be handled in accordance with Clause 3 and Clause 4 Article 34 of this Decree.

Article 36. Handling, repair and recall of defective medical devices [335]

1. Methods for handling defective medical devices[336]:

a) Providing instructions on how to fix the defect;

b) Fixing the defect in medical devices [337];

c) Replace the defective medical device[338] with an equivalent medical device[339];

d) Recall the defective medical device for re-export or destruction;

2. Methods for recall of the defective medical devices[340]:

a) Voluntary recall by the registration number holder;

b) Mandatory recall in the cases specified in Article 39 of this Decree.

Article 37. Handling of medical devices [341] in case product owner [342]  or registration number holder of medical device [343] no longer carries out manufacturing or goes bankrupt or is dissolved

1. If the product owner[345] declares termination of manufacturing of a medical device[344] which has been granted the registration number or goes bankrupt or is dissolved, the medical device may be placed on the market for a maximum period of 24 months from the date of the product owner’s[346] declaration of manufacturing termination, bankruptcy or dissolution provided that the registration number holder in Vietnam provides a commitment to provide warranty or materials for replacing or serving the use of the medical device[347] for 08 years, unless the registration number holder is a permanent representative office in Vietnam of a foreign trader that is the product owner[348].

2. If the registration number holder[350] of a medical device[349] goes bankrupt or is dissolved, the medical device may be placed on the market for a maximum period of 24 months from the date of the registration number holder’s[351] declaration of bankruptcy or dissolution provided that the distributor provides a commitment to provide warranty or materials for replacing or serving the use of the medical device[352] for 08 years.

3. The registration number holder or distributor shall send their commitments to the Ministry of Health through the Portal on management of medical devices[353] within 60 days from the day on which the product owner[354] or registration number holder of the medical device[355] declares termination of manufacturing or goes bankrupt or is dissolved.

4. The commitment shall include the following documents:

a) The written commitment to provide warranty and materials serving the use of the medical devices[356];

b) The list of medical devices[357] of which it is the registration number holder but the product owner[358] or registration number holder of the medical device [359] declares termination of manufacturing or goes bankrupt or is dissolved.

5. Within 15 working days from the receipt of the commitment as prescribed in Clause 4 of this Article, the Ministry of Health shall give a written response indicating whether or not the medical device[360] may be placed on the market. If the placement of the medical device on the market is not approved, reasons for such refusal shall be given.

6. [361] If the Ministry of Health of Vietnam has given a written response indicating its refusal to permit the placement of the medical device[362] on the market as prescribed in Clause 5 of this Article, the registration number holder or distributor shall conduct the recall of medical devices[363] placed on the market, except medical devices[364] sold to health facilities or users.

Section 5 [365]. REVOCATION OF REGISTRATION NUMBER OF MEDICAL DEVICES [366] AND DISPOSAL OF MEDICAL DEVICES[367] WHOSE REGISTRATION NUMBER IS REVOKED

Article 38. Cases where registration number is revoked

1. The application for registration number is forged.

2. In case where 03 batches of the medical device[368] are recalled during the validity of the registration number, except registration number holder’s voluntary recall of the medical device.

3. The applicant deliberately alters or erasures the contents of the registration number.

4. The registration number holder shuts down or is no longer authorized by the product owner[369] and no substitute is appointed, except the case specified in Article 37 of this Decree.

5. Quality of the medical device[370] placed on the market is not consistent with the registered quality.

6. The registration number has been issued against regulations of this Decree.

7. The registration number holder or distributor has not provided commitment for the medical device[371] as prescribed in Clause 1 and Clause 2 Article 37 of this Decree.

8. The period for placing the medical device[372] on the market expires as prescribed in Clause 1 or Clause 2 Article 37 of this Decree.

9. The medical device[373] has been manufactured at a factory that fails to satisfy eligibility requirements laid down in this Decree.

10. The registration number holder fails to comply with the provision of Point k Clause 3 Article 74 of this Decree, except the case specified in Article 37 of this Decree.

11. The application for declaration or issuance of registration number is found to not comply with regulations of this Decree.

12. The medical device[374] has been classified against regulations on classification of medical devices[375].

13. The registration number holder applies for voluntary revocation of the registration number.

14. [376] A document included in the application for registration of medical device[377] is found by a competent authority to be unconformable with regulations of law.

Article 39. Procedures for revocation of registration number

1. If any of the cases specified in Clauses 1 through 12 Article 38 of this Decree is discovered during the inspection, the inspecting agency shall prepare and send an inspection record to the Ministry of Health or Provincial Department of Health that has issued the registration number (hereinafter referred to as “registration number issuer”).

2. Within 05 working days from the receipt of the inspection record specified in Clause 1 of this Article, the registration number issuer shall consider issuing a decision to revoke registration number under their management.

3. After issuing the decision to revoke registration number, the registration number issuer shall:

a) Publish the decision to revoke registration number on its website and send it to the registration number holder, Ministry of Health, Provincial Departments of Health and customs authorities.

b) Remove all information related to the medical device[378] published on its website.

4. When receiving the decision to revoke registration number of the registration number issuer, Provincial Departments of Health shall publish the full text of the decision on their websites and direct specialized agencies to supervise the recall of relevant medical devices[379].

5. Where the registration number holder applies for revocation of registration number, an application indicating reasons for such revocation shall be submitted to the registration number issuer. After receiving the application, the registration number issuer shall carry out procedures for revocation of registration number as prescribed in Clause 2 and Clause 3 of this Article.

6. [380] Within 05 working days from the receipt of the written conclusion from a competent authority as prescribed in Clause 14 Article 38 of this Decree, the registration number issuer shall consider issuing a document on revocation of registration number under its jurisdiction.

Upon receipt of the document on revocation of registration number, relevant competent authorities shall follow the procedures in Clauses 3 and 4 of this Article.

Article 39a. Disposal of medical devices [381] whose registration is revoked [382]

1. Medical devices[383] which have been sold to health facilities or users shall still be used until they are liquidated as prescribed by law or they reach their expiry date, except medical devices[384] containing defects which are harmful to user’s health but cannot be repaired as prescribed in Clause 4 Article 34 of this Decree.

2. Where a medical device[385] whose registration number is revoked is not yet sold to users or health facilities, the registration number holder shall suspend the placement of that medical device[386] on the market and implement measures for recalling it[387].

Chapter VI

MANAGEMENT OF TRADING OF MEDICAL DEVICES [388]

Section 1. ELIGIBILITY REQUIREMENTS TO BE SATISFIED BY TRADING ESTABLISHMENTS [389]

Article 40. Requirements for trading in Class-B, C, D medical devices [390]

To be permitted to trade in Class-B, C, D medical devices, an establishment must:
1. Have at least 01 employee who has an associate degree, or higher, in technology, medicine, pharmacy, chemistry, biology or medical devices, or an associate degree, or higher, suitable for the medical devices[391] sold by the establishment.

2. Have warehouse and vehicles that satisfy the following requirements:

a) The warehouse must:

- have an area suitable for the categories and quantities of medical devices[392] stored therein;

- be well ventilated, dry and clean, and separated from sources of pollution;

- satisfy storage requirements for the medical devices[393] laid down in their user manuals.

b) The vehicle for delivery of medical devices[394] are suitable for the medical devices[395] delivered;

A trading establishment that does not have a warehouse or vehicle for storing or transporting medical devices[396] may sign a contract with a qualified medical device[397] storage or transport service provider.

3. Requirements for trading in medical devices[398] that contain narcotic substances or precursors:

a) The chief technician has a bachelor’s degree in medical devices[399], health, pharmacy, chemistry or biology;

b) The trading establishment must have a warehouse that satisfies the requirements laid down in Article 7 of the Government’s Decree No. 80/2001/ND-CP dated November 05, 2001.

c) There is an inventory monitoring system for medical devices[400] containing narcotic substances and precursors.

Article 41. Application and procedures for declaration of eligibility for medical device[401] trading

1. An application for declaration of eligibility for medical device[402] trading consists of:

a) The declaration of eligibility for medical device[403] trading;

b) A personnel list;

c) Documents proving that the warehouse and vehicle for storing and transporting medical devices[404] satisfy the requirements specified in Clause 2 Article 40 of this Decree. These documents must be certified by the applicant;

d) Documents proving that the warehouse and inventory monitoring system for medical devices[405] containing narcotic substances and precursors comply with the requirements laid down in Clause 3 Article 40 of this Decree. Such documents shall be certified by the establishment making declaration of eligibility for trading in medical devices[406] that contain narcotic substances and precursors.

2. Procedures for declaration of eligibility for medical device trading:

a) Before trading medical devices[407], the trading establishment shall submit an application for declaration of eligibility for medical device[408]  trading to the Department of Health of province where the trading establishment is located.

b) After receiving the application (including the application fee payment as prescribed by the Ministry of Finance), the Department of Health of province where the medical device[409] trading establishment is headquartered shall publish on the Portal on management of medical devices[410] all information about and application for declaration of eligibility for medical device[411] trading.

c) During its operation, the trading establishment shall prepare a notice of changes which is accompanied by supporting documents for such changes, and update such documents to its application for declaration of eligibility for medical device trading published on the Portal on management of medical devices[412] within 03 working days from the occurrence of such changes.

Article 42. Trading in medical devices [413] without satisfying eligibility requirements and following procedures for declaration of eligibility for medical device trading

1. Class-B, C, D medical devices[414] included in the list of medical devices[415] promulgated by the Minister of Health shall be traded as normal goods.

2. Establishments trading in the medical devices[416] prescribed in Clause 1 of this Article shall not be required to satisfy eligibility requirements laid down in Article 40 of this Decree and follow procedures for declaration of eligibility for medical device trading as prescribed in Article 41 of this Article but must still meet the storage and transport requirements as prescribed by product owners[417].

Section 2. MANAGEMENT OF PRICES OF MEDICAL DEVICES [418]

Article 43. Rules for state management of medical device [419] prices

1. Manage prices of medical devices[420] according to the market mechanism; respect the right to pricing and price competition of entities trading in medical devices[421] in accordance with regulations of law.

2. Ensure transparency of prices of medical devices[422] placed on the market.

3. Protect lawful rights and interests of medical device traders and users, and of the State.

4. Take measures for ensuring management of medical device[423] prices appropriate for socio-economic development in each period.

Article 44. Posting prices of medical devices [424] [425]

1. Manufacturers and traders of medical devices[426] shall post prices of medical devices[427] at the locations prescribed in Article 17 of the Government’s Decree No. 177/2013/ND-CP dated November 14, 2013 or on the web portal of the Ministry of Health of Vietnam.

2. The following information must be provided when posting prices of medical devices[428] on the web portal of the Ministry of Health of Vietnam:

a) Name and category of the medical device[429];

b) Manufacturer and manufacturing country; product owner and country of product owner;

c) Unit;

d) Configurations, technical functions of the medical device[430];

dd) Price of medical device[431].

Article 45. Declaring prices of medical devices [432] [433]

1. Manufacturers and traders of medical devices shall declare prices of their medical devices; contents and procedures for declaring prices of medical devices[434] shall comply with regulations of law on pricing.

2. Based on actual situations and upon the occurrence of considerable changes in prices which affect the supply of medical devices[435], payment capacity of buyers and the health insurance fund, the Minister of Health of Vietnam shall promulgate, update and revise the list of medical devices[436] subject to compulsory price declaration and guidelines thereon.

3. Manufacturers and traders of medical devices[437] shall declare prices of their medical devices[438] in the forms specified in the law on pricing or on the web portal of the Ministry of Health of Vietnam.

Section 3. EXPORT AND IMPORT OF MEDICAL DEVICES [439]

Article 46. Rules for management of import and export of medical devices [440] [441]

1. Importers and exporters of medical devices[442] must satisfy eligibility requirements laid down in the law on import and export and assume responsibility to ensure quality, quantities, categories and intended purposes of their imported/exported medical devices[443].

2. Medical devices[444] that have been granted registration numbers in Vietnam may be exported and imported without limits on quantities and are exempt from approval of the Ministry of Health of Vietnam.

3. The import license is required for the medical devices[445] specified in Clause 1 Article 48 of this Decree which are imported for use in Vietnam.

4. The transport of medical devices[446] other than those specified in Clause 2 and Clause 3 of this Article into Vietnam in other forms shall comply with regulations of the law on foreign trade management.

5. Issuance of CFS for medical devices[447] shall comply with regulations of the law on foreign trade management.

6. Import of used medical devices[448] shall comply with regulations of the law on foreign trade management.

Article 47. Export and import of medical devices [449]

1. Domestic enterprises are encouraged to manufacture medical devices for export.

2. An organization or individual that wishes to import the medical device[450] that has been granted registration number shall:

a) be the registration number holder or be authorized in writing by the registration number holder. When authorizing import of medical devices[451], the registration number holder shall send a copy of the authorization letter to the registration number issuer and the customs authority;

b) have a warehouse and vehicles satisfying the requirements laid down in Clause 2 Article 40 of this Decree or have entered into a contract with a qualified medical device[452] storage and transport service provider;

c) have a warehouse and an inventory monitoring system for medical devices[453] containing narcotic substances and precursors that meet the requirements laid down in Clause 3 Article 40 of this Decree.

3. Procedures for import and export of medical devices[454] shall comply with regulations of the law on customs. Importers of medical devices[455] shall not be required to prove their satisfaction of the requirements laid down in Clause 2 of this Article when following customs procedures.

Article 48. Import license

1. An import license is required in the following cases:

a) Unregistered medical devices[456] are imported to serve scientific research, testing, inspection, experiment, performance evaluation, or instruction for use or modification of medical devices[457] only;

b) Unregistered medical devices[458] are imported to serve epidemic prevention and control, or disaster recovery;

c) Unregistered medical devices[459] are imported as assistance or humanitarian aid; gifts or presents given to health facilities; or to serve display at trade fair, exhibition or product launch event;

d) Unregistered medical devices[460] are imported to serve charitable medical examination and treatment;

dd) Unregistered medical devices[461] are imported for personal treatment of illness, including personalized medical devices[462], or to serve a health facility’s special diagnosis demand;

e) [463] Unregistered medical devices[464] are imported for use in health facilities with ODA funding and concessional loans or grants other than ODA grants.

2. An application for the import license consists of:

a) The application form for import license;

b) A synopsis of the technical description of the medical device[465] in Vietnamese;

c) Certificate of conformity with quality control standards of the medical device[466] manufacturer bearing the applicant’s certification;

d) If the medical device is imported to serve research: a certified true copy of the decision to approve the research and documents bearing the applicant’s certification proving that the medical device[467] has been granted marketing authorization by a competent authority;

dd) If the medical device is imported to serve training purposes:  the original copy of the training program and documents bearing the applicant’s certification proving that the medical device[468] has been granted marketing authorization by a competent authority;

e) If the medical device is imported to serve testing, inspection, experiment, or performance evaluation: the certification indicating the quantity of the imported device given by the agency competent to carry out such testing, inspection, experiment, or performance evaluation;

g) If the medical device is imported as aid: a copy of the decision to approve the aid and documents bearing the applicant’s certification proving that the medical device[469] has been granted marketing authorization by a competent authority;

h) If the medical device is imported as gift or present given to a health facility:  the original copy of the training program and documents bearing the applicant’s certification proving that the medical device[470] has been granted marketing authorization by a competent authority;

i) If the medical device is imported to serve charitable medical examination and treatment: documents bearing the applicant’s certification proving that the medical device[471] has been granted marketing authorization by a competent authority;

k) If the medical device is imported to serve a health facility’s special diagnosis demand: documents bearing the applicant’s certification proving that the medical device[472] has been granted marketing authorization by a competent authority;

l) If the medical device is imported to serve personal treatment of illness, including personalized medical devices[473]: a copy of the physician’s prescription which is consistent with the applicant’s illness;

m) If the medical device is imported to serve a trade fair, exhibition, display or product launch event: copies of documents on the program, invitation letter and service contract;

n) If the medical device is imported to serve the purposes of epidemic prevention and control or disaster recovery, the following documents are required:

- A competent authority’s approval for import of the medical device to serve epidemic prevention and control or disaster recovery;

- Documents bearing the applicant’s certification proving that the medical device[474] has been granted marketing authorization or license for emergency use by a competent authority.

o) [475] In the case specified in Point e Clause 1 of this Article, the application for import license shall include:

- The original copies or certified true copies of the decision to approve the investment guidelines and the investment decision for an investment project or the decision to approve project documents for a project on technical assistance, project costs or grants other than ODA grants, in which the import of medical devices[476] must be indicated;

- The original copy of certified true copy of the contract for supply of medical devices[477] for the project;

- The power of attorney granted by the product owner[478] to the applicant which must be still valid at the date of application submission. Either the document bearing consular legalization or the certified true copy thereof is accepted;

- The certificate of eligibility to provide warranty services granted by the product owner[479], except disposable medical devices[480] defined by product owners[481] or cases where there are documents proving that the medical device is not under warranty. Either the document bearing consular legalization or the certified true copy thereof is accepted;

- dd) The unexpired CFS (for imported medical devices[482]). Either the document bearing consular legalization or the certified true copy thereof is accepted. If the CFS is made neither in English nor in Vietnamese, it shall be translated into Vietnamese. The Vietnamese translation must be certified as prescribed by law.

3. Procedures for processing an application for license to import medical devices[483]:

a) If the application for license to import medical devices[484] is satisfactory, the Ministry of Health shall process it within 15 working days, or 02 working days with respect to an unregistered medical device[485] that is imported to serve epidemic prevention and control or disaster recovery, from the receipt of the adequate and valid application (including the application fee receipt as prescribed by the Ministry of Finance). If an application is refused, a written response indicating reasons for such refusal shall be provided.

The import license shall be sent to the applicant and customs authorities.

b) If the application for license to import medical devices[486] is not satisfactory, the Ministry of Health shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant for license to import medical devices[487] within 10 days, or 02 working days with respect to an unregistered medical device [488] that is imported to serve epidemic prevention and control or disaster recovery, from the receipt of the application.

c) The applicant shall comply with the request for modification and send modified documents to the Ministry of Health as requested.

If the modified application is still unsatisfactory, the Ministry of Health shall continue sending another request for modification of the application to the applicant.

d) If the applicant fails to provide the modified application within 30 days from the receipt of the Ministry of Health’s request for modification, the application shall be rejected.

dd) If the modified application is satisfactory, the Ministry of Health shall issue an import license in accordance with Point a of this Clause. The import license shall be sent to the applicant and customs authorities.

Article 49. Application for Certificate of Free Sale (CFS) for medical devices [489]

1. An application for CFS for a medical device[490] consists of:

a) The application form for CFS.

b) A certified true copy of the unexpired certificate of conformity with quality control standards issued by a conformity assessment body as prescribed by law.

2. Procedures for the CFS shall comply with the Government’s Decree No. 69/2018/ND-CP dated May 15, 2018.

Article 50. Authority and procedures for issuance, re-issuance and revocation of CFS

1. The Minister of Health shall consider issuing, re-issuing and revoking CFS for medical devices[491].

2. Procedures for issuance, re-issuance and revocation of CFS shall comply with the Prime Minister’s regulations on issuance of CFS.

Section 4. RIGHTS AND OBLIGATIONS OF ORGANIZATIONS AND INDIVIDUALS ENGAGED IN TRADING OF MEDICAL DEVICES [492]

Article 51. Rights of medical device [493] trading establishments

1. Request the medical device[494] seller to provide adequate information and documents serving the tracing of origin and warranty on the medical device[495].

2. Request importers, distributors and users to cooperate in recalling and handling defective medical devices[496].

3. Request the registration number holder[497] or warranty center that is recognized by the product owner[498] to provide warranty on the medical device[499].

4. Receive notice of defective medical device[500] from the registration number holder.

5. Exercise other rights as prescribed by laws.

Article 52. Obligations of medical device [501] trading establishments

1. Implement internal control measures to maintain the quality of medical devices[502] as prescribed by the registration number holder.

2. Provide users with adequate and timely information about:

a) Instructions for use of medical devices[503]; conditions for ensuring safety, storage, calibration, inspection and maintenance of medical devices[504];

b) Notices of defective medical devices[505].

3. [506] Post and declare prices of medical devices[507] in accordance with the provisions of this Decree and the law on pricing.

4. Keep documents on management of medical devices[508] and carry out tracing of origin and recall of medical devices[509] in accordance with regulations of this Decree.

5. Promptly notify registration number holders and state regulatory authorities of defective medical devices[510].

6. Comply with regulations of law and inspection decisions issued by competent authorities.

7. Perform other obligations as prescribed by law.

Chapter VII

MEDICAL DEVICE-RELATED SERVICES [511]

Section 1. MEDICAL DEVICE[512] TECHNICAL CONSULTING

Article 53. Eligibility requirements to provide medical device [513] technical consulting

1. Consulting on listing and formulation of technical configuration and functions of medical devices[514] must be provided by individuals who have certificates of completion of training course in technical consulting on medical devices.

2. An individual who provides medical device technical consulting is required to:

a) have a bachelor’s degree, or higher, in technology, medicine or pharmacy;

b) have at least 05 years’ experience of working in medical device [515] technology in health facilities;

c) have been examined and recognized by the training institution to be capable of providing medical device[516] technical consulting according to the training program issued by the Ministry of Health.

3. A consultant shall only provide medical device technical consulting after the Ministry of Health publishes his/her information and application for declaration of eligibility to provide medical device technical consulting as prescribed in Point b Clause 2 Article 54 of this Decree.

Article 54. Application and procedures for declaration of eligibility to provide medical device technical consulting

1. An application for declaration of eligibility to provide medical device technical consulting consists of:

a) The application form for declaration of eligibility to provide medical device technical consulting;

b) Certified true copies of qualifications/certificates as prescribed in Point a and Point c Clause 2 Article 53 of this Decree;

c) Certificate of working period.

2. Procedures for declaration of eligibility to provide medical device technical consulting:

a) Before providing medical device technical consulting, the applicant shall send an application for declaration of eligibility to provide medical device technical consulting to the Ministry of Health.

b) After receiving the application (including application fee receipt as prescribed by the Ministry of Finance), the Ministry of Health shall publish on the Portal on management of medical devices[517] all information about and application for declaration of eligibility to provide medical device[518] technical consulting.

c) During its operation, the consultant shall prepare a notice of changes which is accompanied by supporting documents for such changes, and update such documents to his/her application for declaration of eligibility to provide medical device technical consulting published on the Portal on management of medical devices[519] within 03 working days from the occurrence of such changes.

Section 2. INSPECTION AND CALIBRATION OF MEDICAL DEVICES [520]

Article 55. Rules for inspection and calibration of medical devices [521]

1. Medical devices[522] included in the Minister of Health’s list shall undergo safety and function inspection before use (except for the cases in Article 57 of this Decree), periodically and after overhaul. Inspection of medical devices[523] that are measuring devices or radiation equipment shall be carried out in accordance with Clause 2 of this Article.

2. Medical devices[524] that are measuring devices or radiation equipment shall undergo inspection and calibration in accordance with regulations of laws on measurement and atomic energy.

Article 56. Requirements for provision of medical device [525] inspection services

Facilities and personnel requirements, preparation and submission of application for certificate of registration of medical device inspection services, issuance, revision, re-issuance and revocation of certificate of registration of medical device[526] inspection services are the same as those for provision of conformity assessment services.

Regarding the satisfaction of professional requirements for each medical device[527] inspection field: each inspection process registered must be handled by at least 02 inspectors who have certificates of completion of training program in such process.

Article 57. Exemption from first inspection before putting medical devices [528] into service

A medical device[529] shall be exempted from the first inspection before it is put into service in one of the following cases:

1. The certificate of conformity for the medical device[530] is available.

2. The unregistered medical device[531] is imported to serve scientific research or training in use, maintenance or repair of such medical device[532].

3. The unregistered medical device[533] is imported to serve the importer’s personal treatment of illness or charitable medical examination and treatment or special diagnostic purposes.

4. The unregistered medical device[534] is imported for display at a trade fair, exhibition or product launch event.

Article 58. Handling of unqualified medical devices [535]

1. In case the result of the inspection conducted before putting the medical device[536] into service is not satisfactory:

a) Health facilities shall not receive and use that medical device[537];

b) The inspecting organization shall send a written notice of unsatisfactory inspection result to the Ministry of Health;

c) If 03 medical devices[538] of the same batch fail to meet the safety and function inspection requirements, the Ministry of Health shall request the registration number holders in writing to send reports on the quantity of medical devices being placed on the market and those being used in health facilities.

Registration number holders’ reports and the unsatisfactory inspection result are the basis for the Ministry of Health to decide whether to carry out a re-inspection, the quantity of medical devices that have to undergo re-inspection, or suspend the use of such medical devices[539].

The Ministry of Health shall, based on the re-inspection result, decide whether to carry out another re-inspection, the quantity of medical devices that have to undergo another re-inspection, or request the registration number holders to recall the entire batch of medical devices[540].

In case 03 batches of medical devices[541] are recalled during the validity of the registration number, the registration number of the medical device[542] shall be revoked. Medical devices[543] that have been used by health facilities before the issuance of the decision to revoke the registration number may be used if they pass the inspection.

2. In case the result of a periodic inspection or post-overhaul inspection of a medical device[544] is not satisfactory:

a) Health facilities shall not receive and use that medical device[545];

b) The marking of the previous inspection result shall be removed;

c) Health facilities shall cooperate with the registration number holder in implementing corrective measures and carrying out a re-inspection;

d) Only the medical device that passes the re-inspection may be used.

Chapter VIII

MANAGEMENT OF RAW MATERIALS FOR MANUFACTURE OF MEDICAL DEVICES [546], SUBSTANCES FOR EXTERNAL QUALITY ASSESSMENT THAT CONTAIN NARCOTIC SUBSTANCES AND PRECURSORS

Article 59. Rules for management of raw materials for manufacture of medical devices [547], substances for external quality assessment that contain narcotic substances and precursors

1. Concentration and content of raw materials used for manufacture of medical devices[548] and substances for external quality assessment that contain narcotic substances and precursors must be declared before they are imported to or exported from Vietnam.

2. Customs clearance shall be granted according to declaration number and does not require the import license issued by the Ministry of Health.

Article 60. Application and procedures for declaration of raw materials for manufacture of medical devices [549], substances for external quality assessment that contain narcotic substances and precursors

1. An application for declaration consists of:

a) Declaration form of raw materials for manufacture of medical devices[550], substances for external quality assessment that contain narcotic substances and precursors;

b) Certificate of conformity with quality control standards;

c) Technical documents.

2. Procedures for declaration of concentration/content of narcotic substances and precursors:

a) Before importing raw materials for manufacture of medical devices[551] and substances for external quality assessment that contain narcotic substances and precursors, the importer is required to publish an adequate and valid application for declaration as prescribed in Clause 1 of this Article on the Portal on management of medical devices[552];

b) After receiving an adequate and valid application, the Ministry of Health shall publish on the Portal on management of medical devices[553] all information and the application for declaration of raw materials for manufacture of medical devices[554] and substances for external quality assessment that contain narcotic substances and precursors.

3. Importers/exporters shall reapply for declaration of concentration/content of narcotic substances and precursors whenever there are any changes to their declaration.

Chapter IX

INFORMATION ON MEDICAL DEVICES [555]

Article 61. Information on medical devices [556]

1. Information on a medical device[557] is meant to provide medical practitioners and medical device[558] users with instructions on how to use the medical device[559] reasonably and safely.

2. Information on a medical device[560] must be adequate, objective, accurate, truthful and easily understandable and must not cause misunderstanding.

3. Responsibility to provide information on medical devices[561]:

a) Registration number holders and trading establishments of medical devices[562] shall publish information on levels of risks and other information related to the use of medical devices[563];

b) Health facilities shall internally disseminate information on medical devices[564];

c) Healthcare workers shall disseminate information on levels of risks of Class-C, D medical devices[565] to patients;

d) Authorities in charge of managing medical devices[566] shall make information on medical devices[567] publicly available.

4. Providers of information on medical devices[568] shall assume responsibility for their provided information.

5. The Minister of Health shall organize a medical device[569] information system.

Article 62. Advertising of medical devices [570]

1. Contents of an advertisement for a medical device[571] must be consistent with one of the following documents:

a) The application for declaration of applied standards of Class-A or Class-B medical device[572];

b) The application for registration of Class-C or Class-D medical device[573].

2. An advertisement for a medical device[574] shall, inter alia, have the following information:

a) Name, category, product code, manufacturer and manufacturing country of the medical device[575];

b) Registration number;

c) Functions and uses;

d) Name and address of the registration number holder[576] or the entity authorized by the registration number holder[577];

dd) Warnings for medical device users and storage conditions (if any).

3. Audio or video advertisement of medical devices[578] must contain sufficient information specified in Clause 2 of this Article which must be read or displayed clearly.

4. Before carrying out the advertising of a medical device, the registration number holder[579] or the entity authorized in writing by the registration number holder[580] shall publish on the Portal on management of medical devices[581] planned contents and form of advertising.

5. The registration number holder[582] or the entity authorized in writing by the registration number holder[583] shall assume legal responsibility for the conformity of advertisement contents with the published ones and the application for declaration of applied standards of Class-A or Class-B medical device[584] or the application for registration of Class-C or Class-D medical device[585].

6. Documents or materials that do not contain name of a medical device[586], those that contain name and technical specifications of the medical device[587] but do not contain functions or uses of the medical device, scientific research documents, clinical documents, and documents used for training in use of a medical device shall not be considered as advertising documents.

Chapter X

MANAGEMENT AND USE OF MEDICAL DEVICES [588] BY HEALTH FACILITIES

Article 63. Rules for management and use of medical devices [589]

1. Medical devices[590] must be properly managed and used to ensure their functions as well as economical and effective use.

2. Medical devices[591] shall be stored, maintained and used in accordance with the manufacturer's instructions, and must undergo quality in accordance with regulations of this Decree.

With regard to medical devices[592] subject to strict occupational safety and health requirements, regulations of the Law on occupational safety and health must be strictly complied in addition to regulations on quality control laid down in this Decree.

3. Documents on medical devices[593] must be adequately prepared, managed and retained; medical devices[594] shall be sufficiently and punctually recorded according to their actual state and values in accordance with regulations of laws on accounting and statistics and other relevant laws; funding for performing tasks prescribed in Clause 2 of this Article must be ensured.

4. Health facilities shall bear inspection of regulatory authorities in charge of managing medical devices[595].

Article 64. Management and use of medical devices [596] by State-owned health facilities

In addition to the provisions on management and use of medical devices[597] in Article 63 of this Decree, State-owned health facilities shall manage medical devices[598] according to the following provisions:

1. Investment in, purchase, management and use of medical devices[599] shall comply with regulations of law on management and use of public assets.

2. Use of domestically manufactured medical devices[600] are encouraged.

Article 65. Rights and responsibilities of health facilities for management and use of medical devices [601]

1. Each health facility shall have the following rights:

a) Request the registration number holder or the warranty center recognized by the product owner[602] to provide periodical maintenance of medical device during the warranty period;

b) Request the seller to provide technical documents of the medical device[603];

c) Receive the used medical devices[604] to serve their scientific research and training in use and repair of medical devices[605].

2. Each health facility shall have the following responsibilities:

a) Use and operate the medical device[606] according to instructions of the product owner[607];

b) Carry out periodical maintenance, inspection and calibration according to instructions of the product owner[608] or as prescribed by law;

c) Participate in testing for and assessment of quality of medical devices[609];

d) Submit reports on defective medical devices[610] and provide other information at the request of competent authorities.

Chapter XI

ONLINE DECLARATION AND REGISTRATION

Article 66. Online procedures

1. Declaration of eligibility for manufacture of medical devices[611].

2. Declaration of applied standards for medical devices[612].

3. Application for registration of medical device [613].

4. Declaration of eligibility for medical device[614] trading.

5. Declaration of eligibility to provide medical device[615] technical consulting.

6. Application for certificate of registration of medical device[616] inspection services.

7. Application for license to import medical devices[617].

8. Application for Certificate of Free Sale for domestically manufactured medical device[618].

9. Declaration of prices of medical devices[619].

10. Declaration of concentration/content of raw materials for manufacture of medical devices[620] and substances for external quality assessment that contain narcotic substances and precursors.

11. Declaration of contents and form of advertising of medical device[621].

12. [622] If any document required for completing procedures specified in this Decree must be treated as confidential, it shall be submitted in person and managed in accordance with regulations on management of confidential documents.

Article 67. Requirements for online application

An application for declaration, registration, issuance of license or CFS submitted online (hereinafter referred to as “online application”) shall be considered valid when it meets the following requirements:

1. It contains adequate documents which are prepared according to regulations applicable to a physical application and converted into electronic documents. The name of an electronic document must be relevant to the name of the physical document.

2. Information on declaration, application for registration or license is complete and accurate.

Article 68. Procedures for online declaration

1. The legal representative or his/her authorized person shall declare information, upload electronic documents, make certification using digital signature (if any) and make online application fee payment according to procedures on the Portal on management of medical devices[623].

2. The online application receiving authority shall carry out administrative procedures relevant to the application in accordance with regulations of this Decree.

3. Results of online administrative procedures shall have the same legal effect as those of normal administrative procedures.

Article 69. Retention of online application

1. When submitting an online application, the applicant is required to retain physical documents of that application in accordance with Clause 1 Article 33 of this Decree.

2. In case any physical document included in the application specified in Clause 1 of this Article is lost or damaged, the applicant shall give a written notice to the receiving authority, reprepare that document and notify the receiving authority in writing of their completion of application, and then follow procedures for updating their application after obtaining an approval from the receiving authority.

3. Within 35 days from the receipt of the written notice of the loss of document, if the applicant fails to give a written notice of their completion of application, the receiving authority shall:

a) Remove all information published on the Portal about the medical device[624] manufacturer, medical device[625] trading establishment or consultant on medical device[626] technology, raw materials for manufacture of medical devices[627], substances for external quality assessment that contain narcotic substances and precursors, medical device[628] inspection service provider, and registration number of the medical device;

b) Revoke the registration number and license to import medical devices[629].

4. The applicant shall stop their operation and placement of the relevant medical device[630] on the market from the day on which the information is removed as prescribed in Clause 3 of this Article.

Chapter XII

IMPLEMENTATION ORGANIZATION

Article 70. Responsibilities of Ministry of Health

The Ministry of Health shall assume responsibility before the Government for state management of medical devices[631] and have the following tasks and powers:

1. Submit to the Government or Prime Minister to promulgate, or promulgate within their jurisdiction, legislative documents, national technical regulations, strategies, policies and plans on medical devices[632].

2. Direct and organize the implementation of legislative documents, national technical regulations, strategies, policies and plans on medical devices[633].

3. Organize information dissemination and communication on medical devices[634].

4. Provide training for human resources in the field of medical devices[635].

5. [636] Publish the following information on its web portal:

a) Successful bids for procurement of medical devices[637] by state-owned health facilities nationwide;

b) List of medical devices[638] whose registration number has been revoked;

c) List of organizations/individuals forging applications or committing violations against regulations on management of medical devices[639] set out in this Decree.

6. Decide whether or not to apply provisions of this Decree to products or goods are considered as medical devices[640] in some countries but may not be considered as such medical devices[641] in other countries.

7. [642] Play the leading role and cooperate with regulatory authorities in conducting inspections, settling complaints/denunciations and taking actions against violations related to medical devices[643] in accordance with regulations of this Decree and relevant laws.

8. Update and publish the list of organizations specified in Point a Clause 2 Article 29 of this Decree.

9. Provide detailed regulations on classification of medical devices[644] which must be conformable with ASEAN’s treaties on classification of medical devices[645] to which Vietnam is a signatory; issue practical training programs in classification of medical devices[646].

10. Issue the list of medical devices[647] subject to inspection and class-based inspection procedures for listed medical devices[648].

11. Provide guidelines for preparation of ASEAN Common Submission Dossier Template (CSDT) on medical devices[649].

12. Provide specific provisions on templates used during the implementation of this Decree.

13. [650] Publish and revise the list of medical devices[651] subject to compulsory price declaration to meet management requirements and actual situations.

14. [652] Give guidelines for information on medical devices[653] subject to compulsory price declaration.

Article 71. Responsibilities of Ministry of Science and Technology of Vietnam

1. Issue the list of medical devices[654] which are measuring devices subject to sample approval, inspection and calibration after obtaining the consent from the Ministry of Health.

2. Play the leading role and cooperate with the Ministry of Health in formulating national standards for medical devices[655]; quality inspection of medical devices[656] that are measuring devices and radiation equipment.

Article 72. Responsibilities of Ministry of Finance

1. Provide guidelines for management of public assets that are medical devices[657] by State-owned health facilities after obtaining opinions from the Ministry of Health.

2. Provide specific provisions on management and use of fees and charges in the field of medical devices[658] in accordance with regulations of law on fees and charges.

3. Carry out inspection and take actions against violations in the field of medical devices[659].

Article 73. Responsibilities of provincial People's Committees

1. Bear responsibility to manage activities relating to trading and use of medical devices[660] in the province.

2. Organize information dissemination and communication on medical devices[661] in the province.

3. Provide training for human resources in the field of medical devices[662] in the province.

4. Publish on the Portal of the provincial People's Committee and send to the Ministry of Health the following information:

a) Successful bids for medical devices[663] of State-owned health facilities in the province.

b) List of medical devices[664] whose registration number has been revoked in the province.

5. [665] Take responsibility to organize and process procedures as prescribed in this Decree; organize inspections, settle complaints/denunciations and take actions against violations related to medical devices[666] and prices of medical devices[667] in their provinces in accordance with regulations of law.

Article 74. Responsibilities of medical device traders [668]

1. Medical device[669] traders shall assume responsibility for safety and quality of their medical devices[670].

2. Domestic medical device[671] manufacturers shall assume responsibility to manage quality of medical devices[672] during manufacture, transport and storage of medical devices[673] according to their applications for registration number.

3. Registration number holders shall:

a) Carry out classification of medical devices[674], publish on the Ministry of Health’s Portal and assume legal responsibility for their classification results;

Take correction measures for incorrect classification results that reduce the level of risks of classified medical devices[675] or in term of authority specified in this Decree.

b) Follow procedures for declaration of applied standards or application for registration of medical devices[676] in accordance with this Decree. Assume legal responsibility for the accuracy and truthfulness of their applications for registration number;

c) Establish and maintain operation of warranty centers for medical devices[677] or enter in service contracts with qualified medical device[678] warranty service providers, except disposable medical devices[679] as defined by product owners[680] or cases where there are documents proving that the medical device is not under warranty;

d) Prepare and keep documents on management of medical devices[681] and carry out tracing of origin of medical devices[682] as prescribed in this Decree, except disposable medical devices[683] as defined by product owners[684]; submit reports to police authorities on loss of medical devices or raw materials for manufacture of medical devices that contain narcotic substances and precursors;

dd) Provide adequate and accurate information about the product on the label and user manual of the medical devices[685] in accordance with regulations of law on goods labels and provisions of this Decree;

e) Issue prompt, adequate and accurate warnings about risks to users’ health and the environment; instructions for sellers and users on how to minimize the risks; provide information about requirements for transport, storage and use of medical devices[686];

g) Promptly stop placement on the market, take corrective actions against or recall defective medical devices[687] in accordance with this Decree, and notify relevant parties. Destruction of defective medical devices[688] shall comply with regulations of law on environmental protection and relevant laws; the destruction costs shall be paid by the registration number holder;

h) Comply with regulations of law and inspection decisions issued by competent authorities;

i) Pay compensation for damage caused by defective medical devices[689] as prescribed by law;

k) Ensure that the following documents are effective during the effective period of the registration number:

- The CFS (for imported medical devices[690]);

- The authorization letter, except the case specified in Point a Clause 1 Article 25 of this Decree;

- Certificate of eligibility to provide warranty services, except disposable medical devices[691] defined by product owners[692] or cases where there are documents proving that the medical device is not under warranty.

l) Ensure that the medical devices[693] are only manufactured during the effective period of the certificate of conformity with quality control standards;

m) Take legal responsibility for the legitimacy and accuracy of the documents posted while following the procedures specified in this Decree;

n) Provide every health facility that buys the medical devices[694] with 01 set of quality control documents on such medical devices[695] specified in Clause 4 Article 33 of this Decree;

o) [696] Post and declare prices of medical devices[697] in accordance with the provisions of this Decree and the law on pricing.

p) Perform other obligations as prescribed by law.

4. Traders, exporters, importers and transferors of medical devices[698], raw materials, or substances for external quality assessment that contain narcotic substances and precursors shall submit annual reports to the Ministry of Health and Ministry of Public Security by January 15 of the following year.

5. [699] When submitting applications for handling of procedures as prescribed in this Decree, the applying organizations or individuals shall:

a) assume legal responsibility for the accuracy and legitimacy of documents and materials included in their submitted applications;

b) ensure the consistency and conformity of information on medical device[700] provided in the initially submitted request or application with that provided in documents additionally submitted at the request of competent authorities;

c) ensure the validity of submitted documents during the handling of procedures;

d) retain documents included in their submitted applications.

Chapter XIII

IMPLEMENTATION

Article 75. Effect [701]

1. This Decree comes into force from January 01, 2022.

2. The following Decrees cease to have effect from the effective date of this Decree, including:

a) The Government’s Decree No. 36/2016/ND-CP dated May 15, 2016 on management of medical devices[702];

b) The Government’s Decree No. 169/2018/ND-CP dated December 31, 2018 providing amendments to the Government’s Decree No. 36/2016/ND-CP dated May 15, 2016 on management of medical devices [703];

c) The Government’s Decree No. 03/2020/ND-CP dated January 01, 2020 providing amendments to Article 68 of the Government’s Decree No. 36/2016/ND-CP dated May 15, 2016 on management of medical devices[704], as amended by the Government’s Decree No. 169/2018/ND-CP dated December 31, 2018 providing amendments to the Government’s Decree No. 36/2016/ND-CP dated May 15, 2016 on management of medical devices[705].

3. Article 7 of the Government’s Decree No. 181/2013/ND-CP dated November 14, 2013 shall be abrogated from July 01, 2022.

Article 76. Transition [706]

1. Applications for issuance of registration number which have been submitted according to the provisions of the Government’s Decree No. 36/2016/ND-CP dated May 15, 2016 on management of medical devices[707], as amended in the Government’s Decree No. 169/2018/ND-CP and the Government’s Decree No. 03/2020/ND-CP (hereinafter referred to as “Decree No. 36/2016/ND-CP”) before January 01, 2022 but have not been yet processed by the effective date of this Decree shall be processed as follows:

a) With regard to an application for registration of Class-B medical device[708], the Ministry of Health of Vietnam shall instruct the applicant to review the submitted application and follow procedures for declaration of applied standards in accordance with this Decree without paying additional fee;

b) With regard to an application for registration of Class-C or D medical device[709], the Ministry of Health of Vietnam shall issue registration number according to Article 32 of this Decree if it meets all of the requirements laid down in Clause 3 Article 30 of this Decree;

c) The medical device[710] classification result given by a classification body eligible to classify medical devices[711] before the effective date of this Decree may be used.

2. Regulations on validity of import license; import of medical devices[712] which are not included in the list of medical devices[713] subject to import license requirements:

a) An import license for medical devices[714] other than IVD reagents which is issued within the period from January 01, 2018 to December 31, 2021 shall remain valid until December 31, 2024 inclusively;

b) An import license for medical devices[715] which are IVD reagents which is issued within the period from January 01, 2018 to December 31, 2021 shall remain valid until December 31, 2024 inclusively and impose no limit on import quantities;

c) Holders of import license prescribed in Point a and b of this Clause shall meet relevant eligibility requirements laid down in laws and assume responsibility for quality, quantity, categories and uses of their imported medical devices[716]. The Ministry of Health of Vietnam shall carry out inspections and consider revoking the import license in case of violations against regulations on management of medical devices[717];

d) The import of medical devices[718] which are not subject to import license requirements (except insecticidal and germicidal chemicals and preparations for medical and household use which are used for disinfection of medical devices[719] only) and have been classified as Class-C or D medical devices[720] as published on the web portal of the Ministry of Health of Vietnam may continue until the end of December 31, 2024 without limits on import quantities and requiring the Ministry of Health of Vietnam’s confirmation as medical devices[721], regardless of the time of publishing information on such medical devices on the web portal of the Ministry of Health of Vietnam.

When following import procedures, the importer shall declare the number of document indicating medical device[722] classification results given by itself or by a qualified classification body and shall  assume responsibility for quality, quantity, categories and uses of imported medical devices[723].

Customs authorities shall verify the document indicating medical device[724] classification results declared by the importer on the web portal of the Ministry of Health of Vietnam.

3. Regulations on validity of registration number, certificate of registration of medical device, and number of certificate of registration of medical device:

a) Registration numbers issued according to the Decree No. 36/2016/ND-CP before January 01, 2022 shall remain valid indefinitely;

b) Certificates of registration of domestically manufactured medical devices[725] issued before January 01, 2022 shall remain valid until the expiry dates thereon;

c) Registration numbers of medical devices[726] which are IVD reagents issued within the period from January 01, 2014 to December 31, 2019 shall remain valid until December 31, 2024 inclusively;

d) Registration number issued to medical devices[727] that are IVD reagents within the period from January 01, 2020 to December 31, 2021 shall remain valid until the expiry dates on issued certificates of registration;

dd) Holders of certificate of registration or marketing authorization number prescribed in Points b, c and d of this Clause shall meet relevant eligibility requirements laid down in laws and assume responsibility for quality, quantity, categories and uses of their imported medical devices[728]. The Ministry of Health of Vietnam shall carry out inspections and consider revoking certificates of registration or marketing authorization numbers in case of violations against regulations on management of medical devices[729].

4. Regarding applications for import license for medical devices[730] which have been submitted before January 01, 2022 but have not been yet processed:

The Ministry of Health of Vietnam shall inform and request applicants to complete their applications for registration number according to the provisions of the Decree No. 98/2021/ND-CP and prioritize the processing of such applications. If the applicant wants to obtain an import license according to the submitted application, the Ministry of Health of Vietnam shall issue the import license according to the sequence and procedures in Point c of this Clause if the submitted application is adequate and meets the requirements laid down in Point a or b of this Clause.

a) An application for import license for a medical device[731] included in the list of medical devices subject to import license requirements published by the Minister of Health of Vietnam shall include:

- The application form for import license;

- The unexpired CFS of the category of medical device[732] to be imported (original copy or certified true copy);

- The manufacturer’s unexpired certificate of conformity with ISO 13485 quality control standards (the original or copy bearing the applicant’s certification);

- The unexpired power of attorney given by the product owner[733] to the importer of the medical device[734] (the original or certified true copy);

- Technical file describing the category of imported medical device[735] in Vietnamese (bearing the applicant’s certification);

- Catalogue describing functions and technical specifications of category of imported medical device[736];

- Documents on clinical evaluation and instructions for use of the product owner or manufacturer of medical devices[737] which are invasive devices and instruments in cardiology and cranial nerve.

b) An application for import license for IVD reagents shall include:

- Import order;

- The unexpired CFS (original copy or certified true copy);

- The manufacturer’s unexpired certificate of conformity with ISO 13485 quality control standards (the original or copy bearing the applicant’s certification);

- Standards and methods for quality inspection of medical device[738] (bearing the applicant’s certification);

- Label and user manual in Vietnamese accompanied with original label and user manual (bearing the applicant’s certification).

c) Procedures for granting import license:

- If the application for import license does not need to be modified, the Ministry of Health of Vietnam shall issue the import license. If the application for import license is refused, a written response indicating reasons for refusal shall be given;

- If the application for import license needs to be modified, the Ministry of Health of Vietnam shall send a request to the applicant for modification, in which documents and/or contents to be modified must be indicated;

- Upon receipt of the request for modification, the applicant shall provide the modified application to the Ministry of Health of Vietnam within 60 days from the date of the request and shall also assume responsibility for the consistency of contents of the modified application and those of the previously submitted application.

Over 60 days from the receipt of the request for modification from the Ministry of Health of Vietnam, if the applicant fails to submit the modified application or the application is still unsatisfactory after 03 modification times, the Ministry of Health of Vietnam shall refuse to issue the license to import medical devices[739];

d) An import license issued according to the provisions of this Clause shall be valid until December 31, 2024, inclusively.

5. Regulations on application of ASEAN Common Submission Dossier Template (CSDT) on medical devices[740]: CSDT is compulsory from January 01, 2024.

6. Applications for issuance of registration number submitted before January 01, 2024 as prescribed in Article 30 of this Decree:

a) An application for issuance of registration number shall include the documents specified in Article 30 of this Decree of which CSDT document and report on CSDT document validation in Point c Clause 5 Article 30 of this Decree shall be replaced with the following documents:

- The synopsis of technical description of the medical device[741]: the copy in Vietnamese accompanied by technical documents describing functions and specifications of the medical device[742] issued by the product owner[743] and bearing the applicant’s certification shall be submitted. For in-vitro reagents, calibrators and control materials: the synopsis of technical description in Vietnamese must be accompanied by documents stating materials and safety of the product, manufacturing process, pre-clinical and clinical trial reports including stability reports;

- The user manual for medical device[744]: the copy in Vietnamese bearing the applicant’s certification, accompanied by the original copy in English issued by the product owner[745] in case of imported medical device[746];

- Sample of the label for the medical device[747] sold in Vietnam: The sample label bearing the applicant’s certification is submitted. The sample label must meet requirements laid down in regulations of law on labeling of goods.

b) The receipt and processing of applications for registration of medical devices[748] prescribed in Clauses 1, 2, 3 and 4 Article 30 of this Decree shall comply with the provisions of Article 32 of this Decree.

c) An application for registration of a medical device[749] prescribed in Clause 5 Article 30 of this Decree shall be received and processed as follows:

- If the application does not need to be modified, the Minister of Health of Vietnam shall process the application and consider issuing the registration number within 90 days from the receipt of the adequate and valid application (including application fee receipt as prescribed by the Ministry of Finance). If the application is refused, a written response indicating reasons for refusal shall be provided;

- If the application needs to be modified, the Ministry of Health of Vietnam shall send a request for modification, in which such documents and contents requiring modification must be specified, to the applicant that is required to send the modified application to the Ministry of Health of Vietnam within 60 days from the receipt of the request;

- The applicant shall comply with the request for modification and send the modified application to the Ministry of Health.

If the modified application is still unsatisfactory, the Ministry of Health of Vietnam shall continue sending a request for modification to the applicant to modify the application as prescribed in this Clause.

If the applicant fails to provide the modified application within 90 days from the receipt of the Ministry of Health’s request for modification or the application is still unsatisfactory after 03 modification times, the application shall be rejected.

7. The application of the following regulation: “Prohibition of trading in medical devices[750] before their prices are declared or at prices higher than those available on the Ministry of Health’s web portal at the time of trading” is optional with respect to contract packages for bid opening is made before April 01, 2022.

8. Regarding contract packages for purchase of medical devices[751] for which the supplier selection plans have been approved before the effective date of this Decree but bidding documents have yet to be published or issued, the modification of supplier selection plans in respect of price-related contents shall comply with regulations of law on bidding if it is deemed necessary.

Article 77. Responsibility for guidance and implementation

1. The Minister of Health of Vietnam shall instruct, organize and inspect the implementation of this Decree.

2. Ministers, heads of ministerial agencies, heads of Governmental agencies, Chairpersons of Provincial People’s Committees and relevant authorities, organizations and individuals are responsible for the implementation of this Decree./.

|  | CERTIFIED BY PP. MINISTER DEPUTY MINISTER Do Xuan Tuyen |
| --- | --- |

Form 10 - Application for approval of clinical trial made by organization or individual whose medical device is subject to clinical trial [752]

| [NAME OF APPLICANT] ----- | THE SOCIALIST REPUBLIC OF VIET NAM Independence-Freedom-Happiness ----------------- |
| --- | --- |
| No. ... /…….Re: application for approval of clinical trial for medical device | Hanoi, on ... ... ... ...[date] |

To: The Ministry of Health of Vietnam.

Organization/individual owning the medical device: ……………………………………………

ID number/number of citizen identity card or identity card/personal identification number/ passport number1: ……………………………………………………

Business address: …………………………………………………………………………………

Telephone number: …………………………………… Fax: ……………………………………

Email: ………………………………………………………………………………………………

We hereby submits this letter to the Ministry of Health of Vietnam to apply for an approval of clinical trial for medical device with the details as follows: ……………………………………………………………………………………………….

Clinical trial phase: ………………………………………………………………………………

Name of medical device: ………………………………………………………………………

Type: ……………………………………………………………………………………………

Class of medical device: ………………………………………………………………………

Name of the manufacturer: ……………………………………………………………………

Address of the manufacturer: …………………………………………………………………

Proposed clinical trial institution: ………………………………………….

Address: …………………………………………………………………………………………

Telephone number: …………………………………… Fax: …………………………………

Email: ……………………………………………………………………………………………

Full name of the principal investigator: ………………………………………………………

ID number/number of citizen identity card or identity card/personal identification number/ passport number2: ……………………………………………………

Telephone number: …………………………………… Email: …………………………………

Attached documents: 3

1. ……………………………………………………………………………………………

2. ……………………………………………………………………………………………

3. ……………………………………………………………………………………………

The organization/individual owning the medical device and the principal investigator hereby undertake that there is no conflict of interest between parties participating in the trial, and we shall properly conduct the clinical trial approved by the Ministry of Health of Vietnam and strictly comply with good clinical practices and other regulations on clinical trial for medical devices.

| PRINCIPAL INVESTIGATOR (signature and full name) | REPRESENTATIVE OF ORGANIZATION/INDIVIDUAL OWNING MEDICAL DEVICE (signature, full name and seal) |
| --- | --- |

___________________

1Specify any of ID number/number of citizen identity card or identity card/personal identification number/ unexpired passport number.

2Specify any of ID number/number of citizen identity card or identity card/personal identification number/ unexpired passport number.

3List all documents attached to the application form. All documents must be complete and arranged in the order prescribed in the Decree No. 96/2023/ND-CP.

Form 11 - Application for approval of clinical trial made by clinical trial institution [753]

| [NAME OF APPLICANT] ----- | THE SOCIALIST REPUBLIC OF VIET NAM Independence-Freedom-Happiness ----------------- |
| --- | --- |
| No. … / …….Re: application for approval of clinical trial for medical device | Hanoi, on ... ... ... ...[date] |

To: The Ministry of Health of Vietnam.

Name of the clinical trial institution:

Address:

Telephone number:                                                Fax:

Email:

Full name of the principal investigator:

ID number/number of citizen identity card or identity card/personal identification number/ passport number1: ……………………………………………………

Telephone number:                                                 Email:

Organization/individual owning the medical device:

ID number/number of citizen identity card or identity card/personal identification number/ passport number (for individual)2: ……………………………………………………

Business address:

Telephone number:                                                     Fax:

Email:

We hereby submits this letter to the Ministry of Health of Vietnam to apply for an approval of clinical trial for medical device with the details as follows: ……………………………………………………………………………………………….

- Clinical trial phase:

- Name of medical device:

- Type:

- Class of medical device:

- Name of the manufacturer:

- Address of the manufacturer:

- Name of the product owner:

Attached documents: 3

1. ………………………………………………………………………………………………

2. ………………………………………………………………………………………………

3. ………………………………………………………………………………………………

We kindly request the Ministry of Health of Vietnam to consider and give approval of the clinical trial for medical device mentioned above.

The clinical trial institution and the principal investigator hereby undertake that there is no conflict of interest between parties participating in the trial, and we shall properly conduct the clinical trial approved by the Ministry of Health of Vietnam and strictly comply with good clinical practices and other regulations on clinical trial for medical devices.

| PRINCIPAL INVESTIGATOR (signature and full name) | HEAD OF THE CLINICAL TRIAL INSTITUTION (signature, full name and seal) |
| --- | --- |

_________________

1Specify any of ID number/number of citizen identity card or identity card/personal identification number/ unexpired passport number.

2Specify any of ID number/number of citizen identity card or identity card/personal identification number/ unexpired passport number.

3List all documents attached to the application form. All documents must be complete and arranged in the order prescribed in the Decree No. 96/2023/ND-CP.

Form 12 - Description of clinical trial protocol for medical device [754]

MINISTRY OF HEALTH OF VIETNAM
 Description of clinical trial protocol for medical device

I. General information on the clinical trial for medical device

| 1. Name of the trial |  | 1. Name of the trial2. Code (provided by the supervisory authority) |
| --- | --- | --- |
| 3. Trial period: (From ……. (month/year) to....... (month/year)) |  | 4. Level of supervisory authority3. Trial period: State □ Ministry/Provincial government □ (From ……. (month/year) to....... (month/year)) |
| 5. | FundingFunding |  |
| Total:Total:Total: Funding from state budget: ………………… (clearly indicate source of funding: funding for science and technology, donations, equity, etc.) Funding from state budget: ………………… (clearly indicate source of funding: funding for science and technology, donations, equity, etc.) Funding from state budget: ………………… (clearly indicate source of funding: funding for science and technology, donations, equity, etc.) |  |  |
| 6 | Planned trial phase:Planned trial phase: |  |
| 7 | Principal investigator: Principal investigator: |  |
| Full name:Full name:Full name: Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: |  |  |
| 8 | The clinical trial institution:The clinical trial institution: |  |
| Name of the clinical trial institution:Name of the clinical trial institution:Name of the clinical trial institution: Telephone number: Fax: Email: Address: Telephone number: Fax: Email: Address: Telephone number: Fax: Email: Address: |  |  |
| 9 | Organization/individual whose medical device is subject to the clinical trial (is the one entitled to the copyright for the medical device undergoing the clinical trial and use the clinical trial result for initiating manufacture of the medical device or put the medical device into use or following trial phase).Organization/individual whose medical device is subject to the clinical trial (is the one entitled to the copyright for the medical device undergoing the clinical trial and use the clinical trial result for initiating manufacture of the medical device or put the medical device into use or following trial phase). |  |
| Name of organizationName of organizationName of organization Telephone number: Fax: E-mail: Office’s address: Full name (of individual): Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: Telephone number: Fax: E-mail: Office’s address: Full name (of individual): Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: Telephone number: Fax: E-mail: Office’s address: Full name (of individual): Academic degree/title: Scientific title: Telephone number: (Office)/ (Home) Fax: Mobile: E-mail: Office’s address: Home address: |  |  |

II. Contents of clinical trial for medical device

| 10 | Objectives:Objectives:Objectives:Objectives:Objectives: |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 11 | Situations of domestic and foreign clinical trials for the medical deviceSituations of domestic and foreign clinical trials for the medical deviceSituations of domestic and foreign clinical trials for the medical deviceSituations of domestic and foreign clinical trials for the medical deviceSituations of domestic and foreign clinical trials for the medical device |  |  |  |  |
| □ Theme’s status □ New □ Follow the theme finished in the previous phase□ Theme’s status □ New □ Follow the theme finished in the previous phase□ Theme’s status □ New □ Follow the theme finished in the previous phase□ Theme’s status □ New □ Follow the theme finished in the previous phase□ Theme’s status □ New □ Follow the theme finished in the previous phase□ Theme’s status □ New □ Follow the theme finished in the previous phase |  |  |  |  |  |
| □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Description of the medical device (manufacturing technology, technical functions, quality of product, process of use or operation, evaluation method, potential risks and measures for protecting safety of participants and assessors/investigators and other relevant contents) □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: □ Overview of the clinical trial for medical device: Foreign trial: Domestic trial: |  |  |  |  |  |
| □ List of relevant clinical trials for medical device announced within the last 10 years.□ List of relevant clinical trials for medical device announced within the last 10 years.□ List of relevant clinical trials for medical device announced within the last 10 years.□ List of relevant clinical trials for medical device announced within the last 10 years.□ List of relevant clinical trials for medical device announced within the last 10 years.□ List of relevant clinical trials for medical device announced within the last 10 years. |  |  |  |  |  |
| 12 | Approach, study methods and technique (clearly describe the trial design - approach, methods for sample collection, sample size, standards for selecting trial participants, study method, equipment/techniques used - compared with other similar methods, study indicators, technical means, equipment used for determining study indicators; indicate novelty, originality and creativity of this trial design and method).Approach, study methods and technique (clearly describe the trial design - approach, methods for sample collection, sample size, standards for selecting trial participants, study method, equipment/techniques used - compared with other similar methods, study indicators, technical means, equipment used for determining study indicators; indicate novelty, originality and creativity of this trial design and method).Approach, study methods and technique (clearly describe the trial design - approach, methods for sample collection, sample size, standards for selecting trial participants, study method, equipment/techniques used - compared with other similar methods, study indicators, technical means, equipment used for determining study indicators; indicate novelty, originality and creativity of this trial design and method).Approach, study methods and technique (clearly describe the trial design - approach, methods for sample collection, sample size, standards for selecting trial participants, study method, equipment/techniques used - compared with other similar methods, study indicators, technical means, equipment used for determining study indicators; indicate novelty, originality and creativity of this trial design and method).Approach, study methods and technique (clearly describe the trial design - approach, methods for sample collection, sample size, standards for selecting trial participants, study method, equipment/techniques used - compared with other similar methods, study indicators, technical means, equipment used for determining study indicators; indicate novelty, originality and creativity of this trial design and method). |  |  |  |  |
| 12.1 Trial site:12.1 Trial site:12.1 Trial site:12.1 Trial site:12.1 Trial site:12.1 Trial site: 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.2 Trial period: 12.3 Study method: Describe type of evaluation (randomized, blinded, unblinded), design of evaluation (parallel groups, matched-pair technique), blinding techniques (single-blind and double-blind techniques), method and process of random sampling. 12.4 Trial participants: Describe the trial participants (standards of participants and exclusion of potential participants), standard operating procedures (SOPs) for the selection of trial participants; method, standards and time of dividing trial participants into study groups. 12.5 Sample size: Number of participants required to achieve the evaluation objectives based on statistical calculations. 12.6 Medical device undergoing clinical trial: Brief description of the medical device undergoing the clinical trial, including name, type of the medical device, applied technology, main specifications, clinical indications and uses. Specific information on the batch of medical devices used in the trial: name, manufacturer, batch number, date of manufacture, expiry date, and proofs of quality control. Packing, labeling, storage, preservation, transport, instructions for use and attentions during use of the medical device. Analyses and assessment of level of risks/benefits from use of the medical device, potential adverse events of the medical device and safety measures appropriate to patients/trial participants, and healthcare workers. 12.7 Procedures for use of the medical device (develop standard operating procedure - SOPs): Clearly describe and state the use process (use methods and ways, intended users, etc.), diagnosis and treatment periods; person in charge of performing steps of the procedure; evaluation criteria and indicators. 12.8. Combination therapy: Any other therapies that could have been defined or allowed to be used at the same time. 12.9 Trial data and collection method. 12.10 Recording and reporting adverse events: Methods for recording and reporting adverse reactions or adverse events and compliance-related terms. 12.11 Methods for handling adverse events 12.12 Criteria for excluding participants from the trial: Exclusion criteria for participants and instructions for termination of the entire or part of the clinical trial. 12.13 “Blinding” technique and protection of identity of trial participants: Procedures for maintaining list of participant identifiers, treatment coefficient, list of randomly selected participants and/or case report forms (CRFs). These documents must allow identification of each patient or participant as well as data inspection and acquisition. 12.14 Provisions on code breaking: Information on setting up of trial codes, place for storing lists and who is allowed to perform code breaking, when and how the code is broken in cases of emergency. 12.15 Methods of data processing and assessment of results: Describe the methods used for assessing results (including statistical methods) and reports on patients or participants dropping out of the trial. 12.16 Methods for providing information for participants: Information provided for participants include how they will be provided with information on the trial, when and how their participant information sheet and consent forms will be collected. 12.17 Training for clinical trial staff: Provide training for investigators participating in the clinical trial (including principal investigator, coordinator, investigators, pharmacists, nurses, technicians, etc.), including: basis contents of the trial, information on methods for evaluation, SOPs on management and use of the medical device, management and use of accompanied products in the trial regimen. 12.18 Ethical issues: Ethical considerations and measures related to the trial (including methods and procedures for selecting participants, participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). 12.19 Post-trial healthcare: Healthcare provided after the trial and treatment methods applied after the trial. 12.20. Implementation plan 12.21 Plan for monitoring, supervision and inspection: - Supervision of principal investigator and other investigators - Supervision of sponsors/organization or individual owning the medical device subject to the clinical trial - Supervision and inspection by supervisory authorities and Research Ethics Committee. 12.22. SOPs for the clinical trial Ethical issues in biomedical research: (Including participant information sheet and consent form, and commitments to comply with ethical guidelines for clinical trial). |  |  |  |  |  |
| 13 | Contents of the clinical trial (list and describe all contents of the trial, highlight the contents which are new and suitable for resolution of issues set out and achieving the stated trial objectives, including planned activities of cooperation in transferring trial results to users)Contents of the clinical trial (list and describe all contents of the trial, highlight the contents which are new and suitable for resolution of issues set out and achieving the stated trial objectives, including planned activities of cooperation in transferring trial results to users)Contents of the clinical trial (list and describe all contents of the trial, highlight the contents which are new and suitable for resolution of issues set out and achieving the stated trial objectives, including planned activities of cooperation in transferring trial results to users)Contents of the clinical trial (list and describe all contents of the trial, highlight the contents which are new and suitable for resolution of issues set out and achieving the stated trial objectives, including planned activities of cooperation in transferring trial results to users)Contents of the clinical trial (list and describe all contents of the trial, highlight the contents which are new and suitable for resolution of issues set out and achieving the stated trial objectives, including planned activities of cooperation in transferring trial results to users) |  |  |  |  |
| 13.1. Content 1: 13.1. Content 1: 13.1. Content 1: 13.1. Content 1: 13.1. Content 1: 13.1. Content 1: - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. - Task 1: - Task 2: - …………. 13.2. Content 2: - Task 1: - Task 2: - …………. …………. |  |  |  |  |  |
| 14 | International cooperationInternational cooperationInternational cooperationInternational cooperationInternational cooperation |  |  |  |  |
| Name of partner |  |  |  | Name of partnerName of partnerName of partnerContents of cooperationContents of cooperation |  |
| 15 | Implementation scheduleImplementation scheduleImplementation scheduleImplementation scheduleImplementation schedule |  |  |  |  |
| No. | Main contents and tasks (main evaluation milestones) |  | Main contents and tasks (main evaluation milestones)Expected products | Duration (starting time - completion time) | In charge person/authority |
| 1 | 2 |  | 23 | 4 | 5 |

III. Results of the clinical trial

| 16 | Expected resultsExpected resultsExpected results |  |  |
| --- | --- | --- | --- |
| □ Diagram□ Diagram□ Diagram□ Diagram |  |  |  |
| □ Data sheets□ Data sheets□ Data sheets□ Data sheets |  |  |  |
| □ Analytical reports, conclucions on efficacy, suitability and safety of the medical device□ Analytical reports, conclucions on efficacy, suitability and safety of the medical device□ Analytical reports, conclucions on efficacy, suitability and safety of the medical device□ Analytical reports, conclucions on efficacy, suitability and safety of the medical device |  |  |  |
| □ Forecasts for the next phase of the clinical trial□ Forecasts for the next phase of the clinical trial□ Forecasts for the next phase of the clinical trial□ Forecasts for the next phase of the clinical trial |  |  |  |
| □ User manual of the medical device□ User manual of the medical device□ User manual of the medical device□ User manual of the medical device |  |  |  |
| □ Use process□ Use process□ Use process□ Use process |  |  |  |
| □ Other products□ Other products□ Other products□ Other products |  |  |  |
| 17 | Product requirements Product requirements Product requirements |  |  |
| No. | Product’s name | Scientific requirements | Notes |
| (1) | (2) | (3) | (4) |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 18 | Methods for transfer of the trial resultsMethods for transfer of the trial resultsMethods for transfer of the trial results |  |  |
| 19 | Impacts of the trial results (in addition to those mentioned in section 18) Impacts of the trial results (in addition to those mentioned in section 18) Impacts of the trial results (in addition to those mentioned in section 18) |  |  |
| □ Training for science and technology staff□ Training for science and technology staff□ Training for science and technology staff□ Training for science and technology staff □ For relevant scientific fields: □ For society and economy: □ For relevant scientific fields: □ For society and economy: □ For relevant scientific fields: □ For society and economy: □ For relevant scientific fields: □ For society and economy: |  |  |  |

IV. Organizations/individuals participating in clinical trial for medical device

| 20 | Activities of organizations cooperating in clinical trialActivities of organizations cooperating in clinical trialActivities of organizations cooperating in clinical trialActivities of organizations cooperating in clinical trialActivities of organizations cooperating in clinical trial |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| (specify all cooperating organizations and their in-charge activities) (specify all cooperating organizations and their in-charge activities) (specify all cooperating organizations and their in-charge activities) (specify all cooperating organizations and their in-charge activities) (specify all cooperating organizations and their in-charge activities) (specify all cooperating organizations and their in-charge activities) |  |  |  |  |  |
| No. | Name of organization | Address |  | AddressActivities/ contributions to the trialActivities/ contributions to the trial |  |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| … |  |  |  |  |  |
| 21 | Connection with manufacture and lifeConnection with manufacture and lifeConnection with manufacture and lifeConnection with manufacture and lifeConnection with manufacture and life |  |  |  |  |
| (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) (specify manufacturers or users of trial results who have participated in the trial and their in-charge activities) |  |  |  |  |  |
| 22 | Personnel participating in the trialPersonnel participating in the trialPersonnel participating in the trialPersonnel participating in the trialPersonnel participating in the trial |  |  |  |  |
| (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) (specify information on persons making major contributions of all clinical trial institutions and institutions/organizations cooperating in conducting the clinical trial) |  |  |  |  |  |
| No. | Full name |  | Full nameWorking place |  | Working placePercentage (%) of working period for the trial |
| A | Principal investigator |  | Principal investigator |  |  |
| B 1 2 … | Investigators |  | Investigators |  |  |

V. Funding for the clinical trial and funding sources

(provide detailed explanations according to the Appendix enclosed herewith)

Unit: VND million

| 23 | Funding for the trial sorted by type of expendituresFunding for the trial sorted by type of expendituresFunding for the trial sorted by type of expendituresFunding for the trial sorted by type of expendituresFunding for the trial sorted by type of expendituresFunding for the trial sorted by type of expendituresFunding for the trial sorted by type of expenditures |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| No. | Funding sources | Total | IncludingIncludingIncludingIncludingIncluding |  |  |  |  |
|  |  |  | Piecework-based professional services | Raw materials, energy | Equipment, machinery | Construction, minor repairs | Others |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
|  | Total funding Including: State budget-derived funding for scientific activities Other funding sources (specify) - Equity - Others (mobilized capital, etc.) |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |

| HEAD OF CLINICAL TRIAL INSTITUTION (signature, full name and seal) | ...... [location & date] PRINCIPAL INVESTIGATOR (signature and full name) |
| --- | --- |
| ……….. [location & date] DIRECTOR OF ADMINISTRATION OF SCIENCE, TECHNOLOGY AND TRAINING……….. [location & date] DIRECTOR OF ADMINISTRATION OF SCIENCE, TECHNOLOGY AND TRAINING |  |

ESTIMATED COSTS OF THE CLINICAL TRIAL

Unit: VND million

| No. | Expenditures | Total |  | Total Funding sourceFunding sourceFunding source |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Funding | Percentage (%) |  | State budget-derived funding for scientific activities | Donations | Others |
| 1 | Piecework-based professional services |  |  |  |  |  |  |
| 2 | Raw materials, energy |  |  |  |  |  |  |
| 3 | Specialized equipment and machinery |  |  |  |  |  |  |
| 4 | Construction, minor repairs |  |  |  |  |  |  |
| 5 | Others |  |  |  |  |  |  |
| Total |  | Total |  |  |  |  |  |

Explanations about expenditures 
(VND million)

Item 1. Piecework-based professional services

| No. | Description | Total funding | Funding sourceFunding sourceFunding source |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | State budget-derived funding for scientific activities | Donations | Others |
| Total |  | Total |  |  |  |

Item 2. Raw materials, energy

| No. | Description | Unit | Quantity | Unit price | Amount |  | Funding sourceFunding sourceFunding source |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | State budget-derived funding for scientific activities |  | Donations |  | Others |  |
| 2.1 | Raw materials |  |  |  |  |  |  |  |  |  |  |  |
| 2.2 | Tools, spare parts |  |  |  |  |  |  |  |  |  |  |  |
| 2.3 | Energy and fuels |  |  |  |  |  |  |  |  |  |  |  |
|  | - Coal |  |  |  |  |  |  |  |  |  |  |  |
|  | - Electricity | kW/h |  |  |  |  |  |  |  |  |  |  |
|  | - Oil and gas |  |  |  |  |  |  |  |  |  |  |  |
|  | - Other fuels |  |  |  |  |  |  |  |  |  |  |  |
| 2.4 | Water | m3 |  |  |  |  |  |  |  |  |  |  |
| 2.5 | Books, documents and data |  |  |  |  |  |  |  |  |  |  |  |
| Total |  |  |  |  |  | Total Total Total Total |  |  |  |  |  |  |

Item 3. Specialized equipment and machinery

| No. | Description | Unit | Quantity | Unit price | Amount |  | Funding sourceFunding sourceFunding source |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | State budget-derived funding for scientific activities |  | Donations |  | Others |  |
| 3.1 | Purchase of technology equipment |  |  |  |  |  |  |  |  |  |  |  |
| 3.2 | Purchase of evaluation and measuring equipment |  |  |  |  |  |  |  |  |  |  |  |
| 3.3 | Depreciation |  |  |  |  |  |  |  |  |  |  |  |
| 3.4 | Leased equipment |  |  |  |  |  |  |  |  |  |  |  |
| 3.5 | Transport and installation |  |  |  |  |  |  |  |  |  |  |  |
| Total |  |  |  |  |  | Total Total Total Total |  |  |  |  |  |  |

Item 4. Construction, minor repair

| No. | Description | Funding | Funding sourceFunding sourceFunding source |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | State budget-derived funding for scientific activities* | Donations | Others |
| 4.1 | Construction of …..m2 of workshop/laboratory |  |  |  |  |
| 4.2 | Repair of …..m2 of workshop/laboratory |  |  |  |  |
| 4.3 | Installation of electrical and water systems |  |  |  |  |
| 4.4 | Other expenses |  |  |  |  |
|  | Total |  |  |  |  |

Item 5. Others

| No. | Description | Funding | Funding sourceFunding sourceFunding source |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | State budget-derived funding for scientific activities* | Donations | Others |
| 5.1 | Business trip expenses |  |  |  |  |
| 5.2 | Administration expenses |  |  |  |  |
| 5.3 | Costs of evaluation, inspection and acceptance testing |  |  |  |  |
|  | - Evaluation costs |  |  |  |  |
|  | - Appraisal and approval of documents |  |  |  |  |
|  | - Supervision costs |  |  |  |  |
|  | - Costs of intermediate inspection and acceptance testing |  |  |  |  |
|  | - Costs of internal acceptance testing |  |  |  |  |
|  | - Costs of official acceptance testing |  |  |  |  |
| 5.4 | Others |  |  |  |  |
|  | - Training |  |  |  |  |
|  | - Conferences |  |  |  |  |
|  | - Printing of documents, office stationery |  |  |  |  |
|  | - Translation of documents |  |  |  |  |
|  | …. |  |  |  |  |
| 5.5 | Allowances for investigators |  |  |  |  |
|  | Total |  |  |  |  |

*Note: NSSNKH - State budget-derived funding for scientific activities.

Form 13 - Application for approval of changes in clinical trial for medical device [755]

| [NAME OF APPLICANT] ----- | THE SOCIALIST REPUBLIC OF VIET NAM Independence-Freedom-Happiness ----------------- |
| --- | --- |
| No. …/…..Re: application for approval of changes in clinical trial | Hanoi, on ... ... ... ...[date] |

To: The Ministry of Health of Vietnam

Name of the clinical trial institution: ………………………………………………………

Name of organization/individual that owns the medical device undergoing the clinical trial: ………………………………………

Full name of the principal investigator: …………………………………………………………

Number of ID card or Passport: …………………………………………………………………

Telephone number: ……………………………………… Email: ………………………………

The Ministry of Health of Vietnam has granted approval of the clinical trial for the medical device [name of the trial] under the Decision No.      /QD-BYT dated………..

We hereby submit this letter to apply for approval of the following changes:

| No. | Changes | Explanation on change | List of documents concerning the change |
| --- | --- | --- | --- |
| 1. |  |  |  |
| 2. |  |  |  |
| 3. |  |  |  |

Attached documents: 1

1. ……………………………………………………………………………………………………

2. ……………………………………………………………………………………………………

3. ……………………………………………………………………………………………………

After studying the Decree No. 98/2021/ND-CP on management of medical devices, the Decree No. 96/2023/QD-CP elaborating the Law on Medical Examination and Treatment and relevant regulations, we hereby declare that we will fully comply with relevant legislative documents and professional regulations as well as codes of ethics for clinical trials. We kindly request the Ministry of Health (via the Administration of Science Technology and Training) to consider granting approval of our abovementioned changes.

| PRINCIPAL INVESTIGATOR (signature and full name) | HEAD OF THE CLINICAL TRIAL INSTITUTION (signature, full name and seal) | REPRESENTATIVE OF ORGANIZATION/INDIVIDUAL OWNING MEDICAL DEVICE UNDERGOING CLINICAL TRIAL (signature, full name and seal) |
| --- | --- | --- |

_______________________

1List all documents attached to the application form. All documents must be complete and arranged in the order prescribed in the Decree No. 96/2023/ND-CP.

Form 14 - Application for approval of results of clinical trial for medical device [756]

| [NAME OF APPLICANT] ----- | THE SOCIALIST REPUBLIC OF VIET NAM Independence-Freedom-Happiness ----------------- |
| --- | --- |
| No. …/….Re: application for approval of clinical trial results | Hanoi, on ... ... ... ...[date] |

To: The Ministry of Health of Vietnam.

The clinical trial institution:

Name of organization/individual that owns the medical device undergoing the clinical trial:

Cooperating organization:

Principal investigator:

We kindly request the Ministry of Health of Vietnam to consider giving approval of results of the clinical trial for medical device:

- Name of the clinical trial:

- Code of the clinical trial:

- Trial period:

- Clinical trial phase:

- Name of medical device:

- Type:

- Class of medical device:

- Name of the manufacturer:

- Address of the manufacturer:

- Name of the product owner:

- Address of the product owner:

Attached documents:1

1. ……….……….……….……….……….……….……….……….……….……….……….

2. ……….……….……….……….……….……….……….……….……….……….……….

3. ……….……….……….……….……….……….……….……….……….……….……….

| PRINCIPAL INVESTIGATOR (signature and full name) | HEAD OF THE CLINICAL TRIAL INSTITUTION (signature, full name and seal) | REPRESENTATIVE OF ORGANIZATION/INDIVIDUAL OWNING MEDICAL DEVICE UNDERGOING CLINICAL TRIAL (signature, full name and seal) |
| --- | --- | --- |

____________________________

1List all documents attached to the application form. All documents must be complete and arranged in the order prescribed in the Decree No. ……/2023/ND-CP.

Form 15 - Report on results of clinical trial for medical device [757]

Cover page 1

MINISTRY OF HEALTH OF VIETNAM

REPORT ON RESULTS OF CLINICAL TRIAL FOR MEDICAL DEVICE

Name of the clinical trial:

Classification of risks:

Code of the trial (if any):

Clinical trial phase:

Principal investigator:

The clinical trial institution:

Name of organization/individual that owns the medical device undergoing the clinical trial:

Sponsor:

Level of supervisory authority: Ministry of Health of Vietnam

Trial period: (from month/year…..to month/year…..)

Total funding for the trial: ………………………. VND million

Including: Funding for scientific activities: ………………………. VND million

Other funding sources (if any) ………………………. VND million

Year ……

Heading page

MINISTRY OF HEALTH OF VIETNAM

REPORT ON RESULTS OF CLINICAL TRIAL FOR MEDICAL DEVICE

Name of the clinical trial:

Principal investigator:

The clinical trial institution:

Name of organization/individual that owns the medical device undergoing the clinical trial:

Sponsor:

Level of supervisory authority: Ministry of Health of Vietnam

Code of the trial (if any):

Trial period: (From month/year….. to month/year……)

Total funding for the trial: ………………………. VND million

Including: Funding for scientific activities: ………………………. VND million

Other funding sources (if any) ………………………. VND million

Year …

Page 3

REPORT ON RESULTS OF CLINICAL TRIAL FOR MEDICAL DEVICE

1. Name of the clinical trial:

2. Name of medical device:

3. Principal investigator:

4. The clinical trial institution:

5. Name of organization/individual that owns the medical device undergoing the clinical trial:

6. Clinical trial phase:

7. Supervisory authority: Ministry of Health of Vietnam

8. Clerk of the trial (if any):

9. Trial sites:

(a) Trial site 1

- Location:

- Principal investigator:

(b) Trial site 2

- Location:

- Principal investigator:

(c)

10. Trial period: (From month/year….. to month/year……)

11. Name of supervisor

12. Commitments on compliance with GCP requirements

13. Reporting date

Page 4

ABBREVIATIONS

TABLE OF CONTENTS

Part A. Abstract

Part B. Detailed contents of the report

1. Preface:

1.1. Brief description of relevant domestic and overseas trials

1.2. Urgency of the trial

1.3. Assumptions of the trial

1.4. Objectives of the trial

2. Overview of the trial:

2.1. Relevant trials in foreign countries

2.2. Relevant trials in Vietnam

3. Trial participants and study methods:

3.1. Trial design

3.2. Sampling, sample size and trial participants.

3.3. Study method.

3.3.1. Study indicators.

3.3.2. Methods for determining study indicators.

3.3.3. Specific study tools.

3.4. Data processing method.

4. Results of the clinical trial:

4.1. Brief description of the medical device undergoing the trial, including intended function, type of device, technology, characteristics and method of use.

4.2. Analysis of selected documents and data, advantages and disadvantages.

4.3. Evaluation of potential risks and threats, and appropriate measures for protecting safety of patients and healthcare workers.

4.4. Evaluation of cultural, geographical and demographic factors (e.g. age, ethnic group, sex, etc.).

4.5. Equivalent methods of device use according to the same safety standards and considerations on ethical issues.

4.6. Summary of health status of the trial participants.

5. Discussion:

6. Conclusion and recommendations:

7. References:

8. Appendixes (if any):…..

COMPULSORY CONTENTS OF THE REPORT

1. Preface (Introduction)

2. Objectives of the trial.

3. Plans for conducting the clinical trials

3.1. Trial design and plan

3.2. Discussion on the trial design and selection of control

3.3. Selection of the trial participants (target population) (including standards of participants, and standards for exclusion of participants from treatment or the trial)

3.4. The medical device undergoing the clinical trial, accompanied devices/techniques/methods/treatment regimen/medicines/products

3.5. Description of the method for ensuring data quality

3.6. Statistical method specified in the clinical trial protocol and determination of sample size

3.7. Changes occurring during the conduct of the trial and analysis according to the plan.

4. Trial participants (patients/volunteers)

4.1. Current conditions of the trial participants

4.2. Variations compared to clinical trial protocol

5. Evaluation of the clinical trial results

5.1. Data analysis

Clearly determine patients/participants used in analysis of the trial results and those excluded from the trial and causes therefor.

5.2. Anthropological characteristics and other basic characteristics

Provide a summary table of anthropological characteristics of each patient

5.3. Determination of the suitability of the medical device undergoing the clinical trial

Summarize and analyze any result of evaluation of the suitability of each patient/participant for the medical device undergoing the clinical trial.

5.4. Treatment effectiveness and data sheet of each patient

Analysis of treatment effectiveness.

Analysis/statistics.

Provide each patient’s response data sheet.

Method of use of the medical device undergoing the clinical trial; technical procedures, dosage, usage of accompanied devices/techniques/methods/treatment regimen/medicines/products, and the relationship with treatment response.

Provide data of each patient.

Conclusions on treatment effectiveness.

6. Safety evaluation

Carry out analysis of data on the safety at 3 levels:

- Level of exposure () which needs to be inspected to determine the safety of the trial.

- Common adverse effects (AE) and tests or changes that need to be determined, as well as other factors that may affect the frequency of unwanted effects.

- Serious AEs and AEs that commonly occur in participants/patients who must withdraw from the clinical trial ahead of the required schedule, or participants/patients died, regardless of whether such a serious AE is related to the medical device.

7. Level of exposure

The level of exposure needs to be evaluated on the basis of the number of patients who received treatment with use of the medical device, and implementation period.

8. Adverse effects (AEs)

- Summary of AEs.

- Description of AEs.

- Analysis of AEs.

- List of AEs sorted by patients/participants.

9. Deaths and other serious AEs

- List of deaths and SAEs (serious AEs).

- Report on deaths and SAEs (serious AEs), and other AE symptoms.

- Analysis and discussion on deaths and SAEs (serious AEs).

10. Evaluation of tests

- List test values of each patient/participant (as prescribed in Appendix) and abnormal values.

- Evaluate each test parameter.

11. Vital signs, physiological symptoms and other observations related to safety

Analyze vital signs, physiological symptoms and observable changes.

12. Conclusions on safety

Re-state the level of safety of the medical device, changes caused by dosage (if the medical device has accompanied products), AEs resulting in termination of the use of medical device, required medical interventions or death, etc.

13. Discussion and conclusion

Evaluate the effectiveness and safety of the medical device undergoing the clinical trial, relevance of benefits and risks.

14. Relevant tables, diagrams and charts

15. List of references

16. Appendixes: List appendixes included in the report.

[1] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[2] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[3] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[4] The Government’s Decree No. 07/2023/ND-CP dated March 03, 2023 providing amendments to the Government’s Decree No. 98/2021/ND-CP dated November 08, 2021 prescribing medical device management is promulgated pursuant to:

“The Law on Government Organization dated June 19, 2015; the Law on Amendments to the Law on Government Organization and the Law on Local Government Organization dated November 22, 2019;

And at the request of the Minister of Health of Vietnam;”

The Government’s Decree No. 96/2023/ND-CP dated December 30, 2023 on elaboration of the Law on Medical Examination and Treatment is promulgated pursuant to:

“The Law on Government Organization dated June 19, 2015; the Law on Amendments to the Law on Government Organization and the Law on Local Government Organization dated November 22, 2019;

The Law on Medical Examination and Treatment dated January 09, 2023;

And at the request of the Minister of Health of Vietnam;”

[5] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[6] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[7] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[8] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[9] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[10] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[11] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[12] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[13] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[14] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[15] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[16] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[17] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[18] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[19] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[20] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[21] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[22] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[23] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[24] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[25] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[26] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[27] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[28] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[29] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[30] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[31] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[32] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[33] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[34] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[35] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[36] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[37] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[38] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[39] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[40] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[41] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[42] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[43] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[44] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[45] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[46] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[47] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[48] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[49] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[50] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[51] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[52] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[53] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[54] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[55] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[56] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[57] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[58] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[59] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[60] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[61] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[62] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[63] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[64] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[65] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[66] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[67] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[68] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[69] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[70] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[71] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[72] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[73] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[74] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[75] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[76] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[77] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[78] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[79] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[80] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[81] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[82] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[83] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[84] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[85] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[86] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[87] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[88] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[89] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[90] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[91] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[92] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[93] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[94] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[95] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[96] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[97] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[98] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[99] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[100] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[101] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[102] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[103] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[104] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[105] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[106] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[107] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[108] Heading of this Chapter is amended according to Clause 8 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[109] This Article is amended according to Clause 9 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[110] This Article is amended according to Clause 10 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[111] This Article is amended according to Clause 11 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[112] This Article is amended according to Clause 12 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[113] This Article is amended according to Clause 13 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[114] This Article is amended according to Clause 14 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[115] This Article is amended according to Clause 15 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[116] This Article is amended according to Clause 16 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[117] This Article is amended according to Clause 17 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[118] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[119] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[120] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[121] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[122] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[123] This Clause is amended according to Clause 1 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[124] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[125] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[126] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[127] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[128] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[129] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[130] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[131] This Article is amended according to Clause 2 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[132] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[133] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[134] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[135] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[136] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[137] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[138] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[139] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[140] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[141] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[142] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[143] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[144] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[145] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[146] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[147] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[148] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[149] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[150] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[151] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[152] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[153] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[154] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[155] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[156] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[157] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[158] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[159] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[160] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[161] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[162] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[163] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[164] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[165] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[166] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[167] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[168] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[169] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[170] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[171] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[172] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[173] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[174] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[175] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[176] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[177] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[178] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[179] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[180] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[181] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[182] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[183] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[184] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[185] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[186] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[187] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[188] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[189] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[190] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[191] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[192] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[193] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[194] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[195] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[196] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[197] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[198] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[199] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[200] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[201] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[202] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[203] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[204] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[205] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[206] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[207] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[208] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[209] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[210] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[211] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[212] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[213] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[214] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[215] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[216] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[217] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[218] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[219] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[220] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[221] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[222] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[223] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[224] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[225] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[226] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[227] This Clause is added according to point a Clause 19 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[228] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[229] This Point is amended according to point b Clause 19 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[230] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[231] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[232] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[233] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[234] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[235] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[236] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[237] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[238] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[239] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[240] This Clause is added according to point c Clause 19 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[241] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[242] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[243] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[244] This Clause is added according to Point d Clause 19 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[245] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[246] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[247] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[248] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[249] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[250] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[251] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[252] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[253] This Point is amended according to Clause 3 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[254] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[255] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[256] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[257] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[258] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[259] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[260] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[261] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[262] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[263] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[264] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[265] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[266] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[267] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[268] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[269] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[270] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[271] This Article is added according to Clause 21 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[272] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[273] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[274] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[275] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[276] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[277] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[278] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[279] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[280] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[281] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[282] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[283] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[284] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[285] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[286] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[287] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[288] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[289] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[290] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[291] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[292] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[293] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[294] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[295] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[296] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[297] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[298] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[299] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[300] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[301] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[302] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[303] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[304] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[305] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[306] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[307] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[308] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[309] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[310] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[311] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[312] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[313] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[314] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[315] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[316] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[317] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[318] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[319] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[320] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[321] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[322] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[323] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[324] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[325] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[326] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[327] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[328] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[329] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[330] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[331] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[332] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[333] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[334] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[335] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[336] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[337] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[338] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[339] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[340] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[341] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[342] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[343] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[344] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[345] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[346] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[347] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[348] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[349] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[350] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[351] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[352] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[353] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[354] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[355] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[356] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[357] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[358] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[359] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[360] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[361] This Clause is amended according to Clause 4 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[362] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[363] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[364] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[365] Heading of this Section is amended according to Clause 5 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[366] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[367] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[368] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[369] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[370] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[371] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[372] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[373] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[374] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[375] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[376] This Clause is added according to Clause 6 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from Mach 03, 2023.

[377] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[378] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[379] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[380] This Clause is added according to Clause 7 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from Mach 03, 2023.

[381] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[382] This Article is added according to Clause 8 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from Mach 03, 2023.

[383] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[384] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[385] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[386] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[387] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[388] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[389] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[390] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[391] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[392] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[393] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[394] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[395] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[396] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[397] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[398] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[399] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[400] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[401] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[402] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[403] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[404] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[405] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[406] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[407] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[408] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[409] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[410] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[411] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[412] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[413] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[414] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[415] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[416] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[417] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[418] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[419] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[420] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[421] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[422] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[423] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[424] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[425] This Article is amended according to Clause 9 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[426] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[427] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[428] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[429] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[430] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[431] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[432] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[433] This Article is amended according to Clause 10 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[434] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[435] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[436] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[437] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[438] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[439] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[440] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[441] This Article is amended according to Clause 11 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[442] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[443] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[444] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[445] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[446] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[447] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[448] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[449] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[450] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[451] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[452] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[453] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[454] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[455] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[456] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[457] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[458] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[459] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[460] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[461] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[462] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[463] This Point is amended according to point a Clause 12 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[464] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[465] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[466] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[467] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[468] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[469] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[470] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[471] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[472] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[473] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[474] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[475] This Point is added according to Point b Clause 12 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from 03 March, 2023.

[476] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[477] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[478] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[479] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[480] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[481] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[482] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[483] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[484] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[485] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[486] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[487] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[488] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[489] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[490] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[491] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[492] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[493] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[494] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[495] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[496] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[497] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[498] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[499] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[500] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[501] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[502] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[503] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[504] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[505] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[506] This Clause is amended according to Clause 13 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from 03 March, 2023.

[507] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[508] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[509] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[510] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[511] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[512] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[513] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[514] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[515] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[516] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[517] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[518] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[519] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[520] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[521] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[522] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[523] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[524] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[525] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[526] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[527] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[528] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[529] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[530] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[531] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[532] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[533] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[534] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[535] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[536] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[537] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[538] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[539] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[540] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[541] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[542] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[543] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[544] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[545] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[546] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[547] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[548] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[549] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[550] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[551] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[552] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[553] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

554] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[555] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[556] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[557] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[558] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[559] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[560] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[561] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[562] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[563] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[564] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[565] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[566] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[567] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[568] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[569] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[570] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[571] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[572] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[573] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[574] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[575] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[576] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[577] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[578] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[579] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[580] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[581] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[582] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[583] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[584] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[585] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[586] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[587] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[588] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[589] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[590] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[591] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[592] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[593] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[594] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[595] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[596] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[597] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[598] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[599] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[600] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[601] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[602] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[603] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[604] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[605] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[606] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[607] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[608] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[609] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[610] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[611] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[612] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[613] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[614] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[615] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[616] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[617] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[618] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[619] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[620] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[621] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[622] This Clause is added according to Clause 14 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from Mach 03, 2023.

[623] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[624] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[625] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[626] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[627] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[628] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[629] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[630] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[631] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[632] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[633] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[634] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[635] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[636] This Clause is amended according to Point a Clause 15 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[637] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[638] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[639] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[640] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[641] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[642] This Clause is amended according to Point b Clause 15 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[643] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[644] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[645] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[646] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[647] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[648] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[649] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[650] This Clause is added according to point c Clause 15 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[651] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[652] This Clause is added according to point c Clause 15 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[653] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[654] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[655] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[656] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[657] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[658] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[659] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[660] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[661] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[662] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[663] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[664] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[665] This Clause is amended according to Clause 16 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from 03 March, 2023.

[666] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[667] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[668] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[669] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[670] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[671] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[672] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[673] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[674] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[675] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[676] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[677] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[678] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[679] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[680] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[681] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[682] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[683] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[684] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[685] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[686] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[687] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[688] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[689] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[690] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[691] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[692] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[693] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[694] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[695] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[696] This Point is amended according to point a Clause 17 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[697] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[698] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[699] This Clause is added according to Point b Clause 17 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[700] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[701] Articles 2 and 3 of the Government’s Decree No. 07/2023/ND-CP dated March 03, 2023 providing amendments to the Government’s Decree No. 98/2021/ND-CP dated November 08, 2021 prescribing medical device management, coming into force from March 03, 2023, stipulate that:

“Article 2. Implementation

This Decree comes into force from the date on which it is signed.

Article 3. Responsibility for implementation

1. The Minister of Health of Vietnam shall instruct, organize and inspect the implementation of this Decree.

2. Ministers, heads of ministerial agencies, heads of Governmental agencies, Chairpersons of Provincial People’s Committees and relevant authorities, organizations and individuals are responsible for the implementation of this Decree./.”

Clause 1 and Clause 20 Article 147 and Article 148 of the Government’s Decree No. 96/2023/ND-CP dated December 30, 2023 on elaboration of the Law on Medical Examination and Treatment, coming into force from January 01, 2024, stipulate that:

“Article 147. Effect

1. This Decree comes into force from January 01, 2024.

20. Provisions of clause 5 Article 105 of this Decree and clause 19 of this Article shall apply from January 01, 2026.

Article 148. Responsibility for implementation

Ministers, heads of ministerial agencies, heads of Governmental agencies, Chairpersons of Provincial People’s Committees and relevant authorities, organizations and individuals are responsible for the implementation of this Decree.”

[702] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[703] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[704] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[705] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[706] This Article is amended according to Clause 18 Article 1 of the Decree No. 07/2023/ND-CP, coming into force from March 03, 2023.

[707] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[708] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[709] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[710] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[711] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[712] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[713] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[714] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[715] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[716] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[717] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[718] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[719] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[720] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[721] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[722] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[723] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[724] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[725] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[726] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[727] The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[728] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[729] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[730] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[731] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[732] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[733] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[734] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[735] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[736] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[737] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[738] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[739] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[740] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[741] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[742] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[743] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[744] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[745] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[746] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[747] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[748] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[749] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[750] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[751] 727]; The phrase “trang thiết bị y tế” (“medical device”) is changed into "thiết bị y tế” (without change in their meaning) as prescribed in clause 7 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[752] The application form for approval of clinical trial made by an organization or individual whose medical device is subject to the clinical trial included in the application for approval of clinical trial as prescribed in Article 17 of this Decree is amended according to point a Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[753] The application form for approval of clinical trial made by a clinical trial institution included in the application for approval of clinical trial as prescribed in Article 17 of this Decree is amended according to point b Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[754] Description of clinical trial protocol for medical device included in the application for approval of clinical trial as prescribed in Article 17 of this Decree is amended according to point c Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[755] The application form for approval of changes in clinical trial for medical device included in the application for approval of changes in clinical trial as prescribed in Article 17 of this Decree is amended according to point d Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[756] The application form for approval of results of clinical trial for medical device included in the application for approval of results of clinical trial as prescribed in Article 17 of this Decree is amended according to point dd Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

[757] Report on results of clinical trial for medical device included in the application for approval of results of clinical trial as prescribed in Article 17 of this Decree is amended according to point e Clause 18 Article 147 of the Decree No. 96/2023/ND-CP, coming into force from January 01, 2024.

------------------------------------------------------------------------------------------------------
This translation is made by THƯ VIỆN PHÁP LUẬT, Ho Chi Minh City, Vietnam and for reference purposes only. Its copyright is owned by THƯ VIỆN PHÁP LUẬT and protected under Clause 2, Article 14 of the Law on Intellectual Property.Your comments are always welcomed